## 02 - Guangzhou - Test Matrix
row | col | background_group | room | procedure | test_type | mid | ccp | bug | owner | status | progress | hours
|  |  | COM | Building Pre-Survey | TOC | 15305 |  | 99449 | Shuge Miao | CCB | 99%
|  |  | COM | Building Survey Geotop | TOC | 15306 |  | 99919 | Shuge Miao | CCB | 99%
|  |  | GTR1 | GTR1 - Building VTC | TOC | 27776 |  | 99892 | Haiyuan Pan
|  |  | GTR2 | GTR2 - Building VTC | TOC | 27776 |  | 99893 | Haiyuan Pan
|  |  | GTR3 | GTR3 - Building VTC | TOC | 27776 |  | 99894 | Haiyuan Pan
|  |  | COM | Rigging Procedure: Water Conditioners (Generic) | TOC | 1435 |  | 99920 | Pierre Collin | Closed | 100% | 140
|  |  | COM | P+ Beam Line Equipment Rigging - Generic | TOC | 10383 |  | 99921 | Pierre Collin | Closed | 100% | 60
|  |  | COM | Beam Line Stands Rigging - Generic | TOC | 10385 |  | 99922 | Pierre Collin | QA | 100% | 40
2 | 10 | rigging | COM | Rigging of Beam Line Magnets | TOC | 10388 |  | 99923 | Pierre Collin | Closed | 100% | 40
|  |  | COM | General cooling pipes commissionning procedure | TOC | 11290 |  | 99924 | Pierre Collin | Closed | 100% | 41
|  |  | COM | Cyclotron main coils cooling circuit pipes commissioning | TOC | 11291 |  | 99925 | Pierre Collin | Closed | 100% | 40
|  |  | COM | Cyclo vacuum pumping system cooling pipes commissioning | TOC | 11892 |  | 99926 | Andriy Lytovchenko | Closed | 100% | 32
|  |  | COM | Cyclo main coils cooling circuit pipes tightness check | TOC | 11894 |  | 99927 | Pierre Collin | Closed | 100% | 30
|  |  | COM | Cyclo vacuum system cooling circuit pipes tightness check | TOC | 11895 |  | 99928 | Pierre Collin | Closed | 100% | 20
|  |  | COM | Cyclotron Alignment Procedure | TOC | 15307 |  | 99929 | Pierre Collin | QA | 100% | 12
|  |  | COM | Cooling water & compressed air Installation Procedure | TOC | 20105 |  | 99930 | Pierre Collin | Closed | 100% | 40
|  |  | COM | Computer Network Cabling Installation | TOC | 22103 |  | 99931 | Yang Zhao |  | 0% | 2
|  |  | COM | Cyclo 230 Generic Rigging Procedure | TOC | 40998 |  | 99932 | Pierre Collin | Closed | 100%
|  |  | COM | Electrical Cables Installation and Connection Procedures | TOC | 51093 |  | 99933 | Mathieu Eppe |  | 60% | 704
|  |  | COM | Electrical cabinets rigging procedure | TOC | 53321 |  | 99417 | Pierre Collin |  | 99% | 60
|  |  | COM | Building TSS Interconnection | WI | 68410 |  | 99934 | Shuling Peng |  | 35% | 68
|  |  | COM | TSS installation & servicing guide | WI | 67204 |  | 99935 | Shuling Peng |  | 30% | 40
2 | 12 | rigging | GTR1 | GTR1 - Rigging Procedure: GTR Ceiling Equipment | TOC | 1508 |  | 99895 | Pierre Collin |  | 100% | 2
3 | 11 | rigging | GTR1 | GTR1 - 360 degree Gantry Alignment Procedure | test | 15310 |  | 99896 | Pierre Collin | CCB | 100% | 16
4 | 10 | rigging | GTR1 | GTR1 - Gantry Beam Line Elements Alignment | TOC | 15311 |  | 99897 | Pierre Collin | CCB | 100% | 15
|  |  | GTR1 | GTR1 - Addendum to Alignment Procedure of Nozzle & Gantry Beam Line | TOC | 29600 |  | 99898 | Pierre Collin | QA | 100% | 80
2 | 11 | rigging | GTR1 | GTR1 - 2D Gantry Rigging Procedure | TOC | 35263 |  | 99405 | Pierre Collin | Closed | 100% | 200
4 | 11 | rigging | GTR1 | GTR1 - Gantry PBS Dedicated Nozzle Rigging | test | 41084 |  | 99899 | Pierre Collin | Closed | 100% | 2
2 | 8 | pps-rigging | GTR1 | GTR1 - PPS Leoni : Installation procedure | TOC | 53349 |  | 99900 | Pierre Collin | Closed | 100% | 4
|  |  | GTR1 | GTR1 - Wired Hand-Pendant Validation Tests | UT | 64493 |  | 99901 | Shuling Peng |  | 0%
|  |  | GTR1 | GTR1 - HPv3 Maintenance Manual | WI | 54605 |  | 99902 | Shuling Peng
|  |  | GTR1 | WMEBCU Recurrent Integration Test | TP | 72997 |  |  | Shuling Peng
2 | 12 | rigging | GTR2 | GTR2 - Rigging Procedure: GTR Ceiling Equipment | TOC | 1508 |  | 99903 | Pierre Collin |  | 50% | 0
3 | 11 | rigging | GTR2 | GTR2 - 360 degree Gantry Alignment Procedure | test | 15310 |  | 99904 | Pierre Collin | CCB | 100% | 20
4 | 10 | rigging | GTR2 | GTR2 - Gantry Beam Line Elements Alignment | TOC | 15311 |  | 99905 | Pierre Collin | CCB | 100% | 15
|  |  | GTR2 | GTR2 - Addendum to Alignment Procedure of Nozzle & Gantry Beam Line | TOC | 29600 |  | 99906 | Pierre Collin | QA | 100% | 4
2 | 11 | rigging | GTR2 | GTR2 - 2D Gantry Rigging Procedure | TOC | 35263 |  | 99408 | Pierre Collin | Closed | 100% | 180
4 | 11 | rigging | GTR2 | GTR2 - Gantry PBS Dedicated Nozzle Rigging | test | 41084 |  | 99907 | Pierre Collin | Closed | 100% | 4
2 | 8 | pps-rigging | GTR2 | GTR2 - PPS Leoni : Installation procedure | TOC | 53349 |  | 99908 | Pierre Collin | Closed | 100% | 3
|  |  | GTR2 | GTR2 - Wired Hand-Pendant Validation Tests | UT | 64493 |  | 99909 | Shuling Peng |  | 0%
|  |  | GTR2 | GTR2 - HPv3 Maintenance Manual | WI | 54605 |  | 99910 | Shuling Peng
|  |  | GTR2 | WMEBCU Recurrent Integration Test | TP | 72997 |  |  | Shuling Peng
2 | 12 | rigging | GTR3 | GTR3 - Rigging Procedure: GTR Ceiling Equipment | TOC | 1508 |  | 99911 | Pierre Collin |  | 50% | 0
3 | 11 | rigging | GTR3 | GTR3 - 360 degree Gantry Alignment Procedure | test | 15310 |  | 99912 | Pierre Collin | CCB | 100% | 16
4 | 10 | rigging | GTR3 | GTR3 - Gantry Beam Line Elements Alignment | TOC | 15311 |  | 99913 | Pierre Collin | CCB | 100% | 15
|  |  | GTR3 | GTR3 - Addendum to Alignment Procedure of Nozzle & Gantry Beam Line | TOC | 29600 |  | 99914 | Pierre Collin | QA | 100% | 16
2 | 11 | rigging | GTR3 | GTR3 - 2D Gantry Rigging Procedure | TOC | 35263 |  | 99409 | Pierre Collin | Closed | 95% | 190
4 | 11 | rigging | GTR3 | GTR3 - Gantry PBS Dedicated Nozzle Rigging | test | 41084 |  | 99915 | Pierre Collin | QA | 100% | 5
2 | 8 | pps-rigging | GTR3 | GTR3 - PPS Leoni : Installation procedure | TOC | 53349 |  | 99916 | Pierre Collin | Closed | 100% | 5
|  |  | GTR3 | GTR3 - Wired Hand-Pendant Validation Tests | UT | 64493 |  | 99917 | Shuling Peng |  | 90% | 16
|  |  | GTR3 | GTR3 - HPv3 Maintenance Manual | WI | 54605 |  | 99918 | Shuling Peng |  | 65% | 40
|  |  | GTR3 | WMEBCU Recurrent Integration Test | TP | 72997 |  |  | Shuling Peng |  | 15%
|  |  | COM | SRCUB_Cabinet_Signal_Test | UT | 55828 | Y | 99453 | Shuling Peng |  | 95% | 128
|  |  | COM | SRCU Configuration Site Check integration Test | VT | 61545 | Y | 99454 | Shuling Peng | QA | 100% | 8
|  |  | COM | P+ TSS validation procedure for CBTV PlatformD&E | VT | 61791 | Y | 99455 | Shuling Peng |  | 30% | 4
|  |  | COM | Radial Probe_ integration test procedure | IT | 1758 |  | 99457 | Shuge Miao | QA | 100% | 10
|  |  | COM | C230 source vacuum Integration Test procedure | IT | 1759 |  | 99458 | Shuge Miao | QA | 100% | 5
|  |  | COM | Cyclotron deflector integration test procedure | IT | 1760 |  | 99459 | Mohan Naikare | QA | 100% | 18
|  |  | COM | Cyclone 230 water cooling integration test | IT | 1761 |  | 99460 | Mohan Naikare | Closed | 100% | 5
|  |  | COM | Cyclo and Beam Line water conditioner unit test procedure | UT | 3514 |  | 99461 | Haiyuan Pan |  | 90% | 40
|  |  | COM | Cyclotron Water manifold unit test procedure | UT | 3515 |  | 99462 | Jean-Michel Denis | Closed | 100%
|  |  | COM | ODP/CRYOs Water Cooling Circuit Unit Test procedure | UT | 3516 |  | 99465 | Andriy Lytovchenko | Closed | 100% | 20
|  |  | COM | Cyclo primary vacuum circuit unit test procedure | UT | 3517 |  | 99466 | Jean-Michel Denis | Closed | 100% | 59
|  |  | COM | Cyclotron air manifold unit test procedure | UT | 3518 |  | 99467 | Christophe COMBLEZ | Closed | 100% | 0.5
|  |  | COM | Cyclotron Main Coils Water Manifold unit test procedure | UT | 3520 |  | 99468 | Shuge Miao | Closed | 100% | 2
|  |  | COM | Cyclotron RF Cavity Resonance Frequency Verification and Adjustment Procedure | UT | 3830 |  | 99469 | Jean-Michel Denis | Closed | 100% | 5
|  |  | COM | 235 Mev Cyclotron RF Cavity Adjustment Procedure for Coupler and Pick-up Antennae | UT | 3831 |  | 99470 | Jean-Michel Denis | Closed | 100% | 8
|  |  | COM | Test Procedure of the C235 amplifier chain | IT | 3833 |  | 99471 | Pascal PELERIN | Closed | 100% | 18
|  |  | COM | C230 vacuum with ODP integration test procedure | IT | 6173 |  | 99472 | Shuge Miao | QA | 100% | 24
|  |  | COM | Cyclo_Beam Lines_Gantries Power Distribution_TPO | UT | 7187 |  | 99473 | Shuling Peng |  | 90% | 8
|  |  | COM | Generic Uninterrupted Power Supply Distribution_TPO.doc | UT | 7188 |  | 99474 | Shuling Peng | Eng-fix | 50% | 2
|  |  | COM | Cyclo Coils integration test procedure | IT | 7493 |  | 99475 | Jean-Michel Denis | QA | 100% | 40
|  |  | COM | Cyclo RF Polarization integration test procedure | IT | 7494 |  | 99476 | Pascal PELERIN | Closed | 100% | 7
|  |  | COM | Cyclo Injection integration test procedure | IT | 7495 |  | 99477 | Mike Welch | Closed | 100% | 20
|  |  | COM | Cyclo RF Power integration test procedure | IT | 7497 |  | 99478 | Pascal PELERIN | Closed | 100%
|  |  | COM | Cyclo Beam Production_on site validation test procedure | VT | 7525 | Y | 99479 | Shuge Miao |  | 95% | 40
|  |  | COM | C230 Main coils water conditioner unit test procedure | UT | 7692 | Y | 99480 | Rene BREEMANS | QA | 100% | 10
|  |  | COM | C230 Yoke Lifting System Unit Test | UT | 37919 | Y | 99481 | Shuge Miao | QA | 100% | 70
|  |  | COM | LLRF Tuning | IT | 42651 |  | 99482 | Shuge Miao | Eng-fix | 99% | 16
|  |  | COM | C230 water leak detection and condensation Test Procedure | IT | 47170 |  | 99483 | Shuge Miao |  | 0%
|  |  | COM | C230_Yoke_Heater_Unit_test | UT | 48694 |  | 99484 | Mike Welch | Closed | 100% | 19
|  |  | COM | BCREU On-Site Validation | VT | 20838 | Y | 99485 | Shuge Miao |  | 95%
|  |  | COM | IC cyclo gains calibration procedure | UT | 54173 |  | 99486 | Shuge Miao |  | 95% | 30
|  |  | COM | R2VCU hardware installation work instruction | WI | 75556 |  |  | Chenchen Yuan |  | 99% | 16
|  |  | COM | Group3 on site validation procedure | VT | 7198 |  | 99493 | Chenchen Yuan | QA | 100% | 16
|  |  | COM | Beam Line Vacuum Circuit on site unit test | UT | 7706 |  | 99494 | Chenchen Yuan |  | 50% | 6
|  |  | COM | Beam Line & PS Room Water Cooling On Site Unit Test Proc. | UT | 7761 |  | 99509 | Yeehim Him |  | 50% | 16
|  |  | COM | BPM, Beam Stop & Vacuum Isolation Valve on site unit test | UT | 16066 |  | 99510 | Chenchen Yuan |  | 99% | 10
|  |  | COM | Degrador integration test procedure | IT | 27887 |  | 99511 | Chenchen Yuan |  | 99% | 10
|  |  | COM | SUT Beam Line Electrometers | UT | 27975 |  | 99512 | Andriy Lytovchenko | Closed | 100% | 113
|  |  | COM | TDSPS On Site Test | UT | 29143 |  | 99513 | Yeehim Him |  | 60% | 9
|  |  | COM | Site integration test for MPSS | IT | 33614 |  | 99514 | Yeehim Him |  | 90% | 48
|  |  | COM | Renovated slit site integration test procedure | IT | 38241 |  | 99515 | Chenchen Yuan |  | 99% | 12
|  |  | COM | BLCU-B - Recurrent TPO | IT | 56455 |  | 99516 | Chenchen Yuan |  | 70% | 28
|  |  | COM | UT Startup BDCU C230 | UT | 60609 |  | 99517 | Yang Zhao |  | 50% | 36
|  |  | COM | BeamLine Integration Test (sites using BLCU-B) | IT | 64140 |  | 99518 | Chenchen Yuan |  | 10% | 2
|  |  | COM | PBSv11 hardware installation guide and FRU replacement proce | WI | 60486 |  | 99519 | Chenchen Yuan |  | 80% | 24
|  |  | COM | WI-MPSS istallation | WI | 66696 |  | 99520 | Chenchen Yuan |  | 99% | 26
|  |  | COM | WI SMPS Installation PPLUS | WI | 67096 |  | 99521 | Chenchen Yuan |  | 50% | 10
|  |  | COM | Beam on degrader focalization | VT | 3233 |  | 99522 | Yeehim Him | QA | 100% | 6
|  |  | COM | Degrader calibration on site validation test procedure | VT | 3242 | Y | 99523 | Chenchen Yuan | QA | 100% | 28
|  |  | COM | Alignment of the cyclotron beam on the ESS axis | VT | 3250 |  | 99524 | Yeehim Him | QA | 100% | 10
|  |  | COM | ESS manual tuning on site validation test procedure | VT | 3252 |  | 99525 | Chenchen Yuan | QA | 100% | 20
|  |  | COM | ESS Range Accuracy and Beam quality Test | VT | 7200 | Y | 99526 | Chenchen Yuan | QA | 100% | 16
|  |  | COM | R2VCU Recurrent integration test | TP | 65063 |  |  | Chenchen Yuan
|  |  | COM | Creation of the Lynx deformations maps | UT | 44448 |  | 99527 | Chenchen Yuan
|  |  | COM | Generation of telescope optical solution | VT | 15276 |  | 99530 | Chenchen Yuan
|  |  | COM | Calibration of the LLRF settings for dark current |  | 19373 | Y | 99531 | Shuge Miao
|  |  | COM | Non Ionization Electromagnetic Emission Level Measurements |  | 4438 | Y | 99532 | Shuling Peng
|  |  | COM | IEC Compliance test of Electrical Installation on PT site |  | 5525 | Y | 99533 | Shuling Peng
|  |  | COM | Generic Uninterrupted Power Supply Capacity TPO | UT | 7657 | Y | 99534 | Shuling Peng
|  |  | COM | Installation and check of safety labels in TC |  | 24717 | Y | 99535 | Shuling Peng
|  |  | COM | IEC 60601-2-64 Site Test Range & Beam Line interlocks P-PLUS | TP | 61281 |  |  | Chenchen Yuan
|  |  | COM | BMS PBS Test case: data recording failure cases (C230) |  | 62554 | Y | 99537 | Chenchen Yuan
|  |  | COM | PRIDEx cloud connector installation | WI | 73288 |  |  | Yang Zhao
|  |  | COM | BIREU configuration site test | VT | 48827 | Y | 99538 | Shuge Miao
|  |  | COM | Safety of clinical runtime environment VTC | VT | 52934 | Y | 99529 | Yang Zhao
|  |  | COM | BAEU configuration site test | VT | 54233 | Y | 99528 | Shuge Miao
|  |  | GTR1 | GTR1 - SRCUC cabinet signals test | UT | 57390 | Y | 99539 | Shuling Peng
|  |  | GTR1 | GTR1 - SRCU Configuration Site Check integration Test | VT | 61545 | Y | 99540 | Shuling Peng
|  |  | GTR1 | GTR1 - P+ TSS validation procedure for Treatment Room PlatformD&E | VT | 61792 | Y | 99541 | Shuling Peng
|  |  | GTR1 | GTR1 - Telescope (straight line) Manual Beam Tuning | VT | 7204 |  | 99542 | Chenchen Yuan
|  |  | GTR1 | GTR1 - Preparing ESBTS-Convalgo | VT | 15288 |  | 99543 | Chenchen Yuan
|  |  | GTR1 | GTR1 - TPO for ESS efficiency characterization | VT | 20206 |  | 99544 | Chenchen Yuan
|  |  | GTR1 | GTR1 - PBS beam tuning | VT | 21888 |  | 99545 | Chenchen Yuan
|  |  | GTR1 | GTR1 - Calibration of IC1 tuning gains in PBS | VT | 30685 |  | 99546 | Chenchen Yuan
|  |  | GTR1 | GTR1 - Field regulation and volumetric repainting calibration | VT | 42969 |  | 99547 | Chenchen Yuan
9 | 8 | function_pps | GTR1 | Correction of the offset of joint A6 of PPS LEONI | test | 68811
5 | 9 | rigging | GTR1 | GTR1 - PT project: Gantry Limit Switches Box on site unit test procedure | test | 1561 |  | 99548 |  |  | 5% | 2
12 | 11 | function | GTR1 | GTR1 - Gantry on site integration test procedure | test | 1801 | Y | 99549 |  | Eng-fix
18 | 11 | validation | GTR1 | GTR1 - Gantry on site validation test procedure | test | 2125 | Y | 99550
5 | 11 | rigging | GTR1 | GTR1 - Gantry Brakes Cabinets on site unit test procedure | test | 2515 |  | 99551 |  |  | 99% | 4
5 | 12 | rigging | GTR1 | GTR1 - Gantry Drivers Cabinet on site unit test procedure | test | 5809 |  | 99552 |  |  | 70% | 5
18 | 12 | validation | GTR1 | GTR1 - GRF on-site recurrent unit tests procedure | test | 19666 |  | 99553
19 | 12 | validation | GTR1 | GTR1 - GRF On-site recurrent integration tests procedure | test | 19667 |  | 99554
20 | 12 | validation | GTR1 | GTR1 - GRF on-site recurrent validation tests procedure | test | 19668 | Y | 99555
15 | 12 | function | GTR1 | GTR1 - GRF TOC | TOC | 21058 | Y | 99556
9 | 12 | function | GTR1 | GTR1 - PCUIO unit test | test | 36877 |  | 99557
11 | 7 | function | GTR1 | GTR1 - Dedicated Nozzle Snout Holder (DNSH) IT test procedure | test | 37826 |  | 99558 | Jianlong Wen |  | 1% | 2
15 | 7 | validation | GTR1 | GTR1 - DNSH alignment procedure using PPVS | test | 43837 |  | 99559
4 | 8 | pps-rigging | GTR1 | GTR1 - Leoni PPS: Site Acceptance Test | test | 57306 | Y | 99560
7 | 8 | pps-rigging | GTR1 | GTR1 - Leoni PPS: Accuracy Validation | test | 57307 | Y | 99561
12 | 8 | function_pps | GTR1 | GTR1 - Leoni Robot PPS Robot recurrent test | test | 58339 | Y | 99562
7 | 9 | rigging | GTR1 | GTR1 - Gantry balancing procedure | test | 62212 | Y | 99563 | Jianlong Wen |  | 20% | 1
25 | 7 |  | GTR1 | GTR1 - IEC60601-2-64_SiteTest_PosDevice I&D | test | 64008 |  | 99564
7 | 10 | rigging | GTR1 | GTR1 - Gantry 360 preliminary mechanical test procedure | test | 65492 |  | 99565
10 | 8 | function_pps | GTR1 | GTR1 - Leoni PPS - Update of Couch Definition File | test | 66850 |  | 99566
11 | 8 | function_pps | GTR1 | GTR1 - Leoni Robot PPS couch definition file verification | test | 64474 |  | 99567
22 | 8 |  | GTR1 | GTR1 - PMS SRIT | test | 64181 |  | 99568
4 | 5 | rigging | GTR1 | GTR1 - Flat panels GTR Mechanical structure on site unit test procedure | test | 2017 |  | 99569
6 | 5 | rigging | GTR1 | GTR1 - Flat Panels Motion Electronic Unit GTR on site unit test procedure | test | 2019 |  | 99570
16 | 6 | rigging | GTR1 | GTR1 - Nozzle X Ray tube stray radiation measurement | toc | 8218 | Y | 99571
5 | 3 | rigging | GTR1 | GTR1 - Laser Protection RVT | test | 26775 |  | 99572
18 | 6 | validation | GTR1 | GTR1 - Test of the radiographic PPVS image quality | test | 27834 | Y | 99573
18 | 1 | validation | GTR1 | GTR1 - CBCT Validation procedure | cbct | 32137 | Y | 99574
5 | 1 | rigging | GTR1 | GTR1 - Configuration procedure for the Rotoscan RS4-4 | cbct | 32138 | Y | 99575
22 | 3 |  | GTR1 | GTR1 - ProteusPlus PPVS SW: non-regression at installation | test | 67728 | Y | 99576
17 | 3 | rigging | GTR1 | GTR1 - Calibration dose vs radiographic tube settings | cbct | 34959 |  | 99577
14 | 4 | rigging | GTR1 | GTR1 - Validation of adaPT insight Geometric Calibration | test | 36048 |  | 99578
4 | 3 | rigging | GTR1 | GTR1 - 38541 CBCT IVV Radioscopic Devices Geometrical Check | test | 38541 |  | 99579
|  |  | GTR1 | GTR1 - CBCT: integration test for SRCU – PCU – X-ray Generator | IT | 43710 | Y | 99580
10 | 4 | rigging | GTR1 | GTR1 - EMD generator calibration procedure | test | 54462 |  | 99581
19 | 6 | validation | GTR1 | GTR1 - PPVS software validation - First installation | test | 54534 | Y | 99582
13 | 5 | rigging | GTR1 | GTR1 - Fine Alignment and FP calibration | test | 55357 | Y | 99583
10 | 3 | rigging | GTR1 | GTR1 - Validation of HW x-ray equipment using EMD generators | test | 56498 |  | 99584
5 | 4 | rigging | GTR1 | GTR1 - PPVS EMD installation guidelines | test | 64208 |  | 99585
4 | 4 | rigging | GTR1 | GTR1 - PPVS Hardware Inspection | test | 60180 |  | 99586
15 | 8 | validation_pps | GTR1 | GTR1 - Treatment couch recurrent validation | test | 65546 |  | 99587
|  |  | GTR1 | GTR1 - IEC60601-2-68 PMS Site tests | VT | 69408 |  | 99588
|  |  | GTR1 | GTR1 - IEC60601-2-68-Site-Data | VT | 67876 |  | 99589
|  |  | GTR1 | GTR1 - IEC60601-2-68-Site-GUI | VT | 68250 |  | 99590
|  |  | GTR1 | GTR1 - IEC60601-2-68-Site-Safety | VT | 68251 |  | 99591
|  |  | GTR1 | GTR1 - adaPT insight safety labels | VT | 68891 |  | 99592
14 | 6 | rigging | GTR1 | GTR1 - Check image orientation in AI | test | 64957 |  | 99593
6 | 3 | rigging | GTR1 | GTR1 - Testing/Validation Installation of adaPT insight | test | 66442 |  | 99594
16 | 8 | validation_pps | GTR1 | GTR1 - PPS Real Iso Offset computation | test | 32781 |  | 99595
15 | 5 | rigging | GTR1 | GTR1 - Adaptinsight_geometric_calibration_procedure | test | 36050 |  | 99596
22 | 4 |  | GTR1 | GTR1 - Clinical Installation of adaPT insight | test | 66385 |  | 99597
16 | 3 | rigging | GTR1 | GTR1 - Acquisitions for CBCT calibration | cbct | 66505 |  | 99598
16 | 2 | rigging | GTR1 | GTR1 - Creation of calibration maps for CBCT | cbct | 66519 |  | 99599
16 | 1 | rigging | GTR1 | GTR1 - Calibration of CBCT reconstruction parameters | cbct | 66530 |  | 99600
|  |  | GTR1 | GTR1 - Site integration test LPIC1 | IT | 43133 |  | 99601
24 | 7 |  | GTR1 | GTR1 - IEC60601-2-64 Site Tests Grade B/C: Accessories | test | 61271 |  | 99602
5 | 10 | rigging | GTR1 | GTR1 - GTR Nozzle Alignment Procedure | test | 15312 |  | 99603
|  |  | GTR1 | GTR1 - BDSIOS PBS Site Integration Tests | IT | 19058 |  | 99604
|  |  | GTR1 | GTR1 - SMPS Integration Test - Dekimo | IT | 45642 |  | 99605
|  |  | GTR1 | GTR1 - PBS system feedback delay calibration |  | 19376 |  | 99606
|  |  | GTR1 | GTR1 - Alignment of PBS unscanned beam |  | 21877 |  | 99607
|  |  | GTR1 | GTR1 - SAD in PBS mode |  | 21900 |  | 99608
|  |  | GTR1 | GTR1 - BDS scanning controller/scanning magnet calibration |  | 21910 |  | 99609
|  |  | GTR1 | GTR1 - PBS ICs parameters calibration |  | 22014 |  | 99610
|  |  | GTR1 | GTR1 - PBS Nozzle WET |  | 22019 |  | 99611
|  |  | GTR1 | GTR1 - IC beam size |  | 22020 |  | 99612
|  |  | GTR1 | GTR1 - PBS thresholds and safety check calibration |  | 22086 | Y | 99613
|  |  | GTR1 | GTR1 - PBS - Wiring check | WI | 33114 |  | 99614
|  |  | GTR1 | GTR1 - Fine tune distance from IC to isocenter |  | 42835 |  | 99615
22 | 2 |  | GTR1 | GTR1 - Xray Proton Colinearity for PPVS testing | test | 53920 | Y | 99616
|  |  | xxGTR1 | GTR1 - Validation procedure: target positioning in the gantry room | VT | 6712 | Y | 99617 |  | Removed
|  |  | GTR1 | GTR1 - UBTI_service_IT | IT | 27719 | Y | 99618
24 | 6 |  | GTR1 | GTR1 - Move PPDs into SETUP and TREATMENT | test | 1918 |  | 99619
23 | 11 | tss | GTR1 | GTR1 - Enable switch: emergency stop switches the hand pendant ON | test | 1929 | Y | 99620
23 | 9 | tss | GTR1 | GTR1 - Collision detection | test | 1930 | Y | 99621
23 | 8 | tss | GTR1 | GTR1 - PPD speed control Validation Test Procedure | test | 1931 | Y | 99622
23 | 7 | tss | GTR1 | GTR1 - PROXIMITY validation Tests | test | 1932 | Y | 99623
23 | 6 | tss | GTR1 | GTR1 - PPDs check during treatment | test | 1975 | Y | 99624
|  |  | GTR1 | GTR1 - BMS - PBS - Test Case 1 - TPO |  | 22058 | Y | 99625
|  |  | GTR1 | GTR1 - BMS PBS DCEU Tests |  | 22069 | Y | 99626
|  |  | GTR1 | GTR1 - Safety Check for PBS Scatterer Position |  | 36485 | Y | 99627
|  |  | GTR1 | GTR1 - Beam veto ITC | IT | 41020 | Y | 99628
|  |  | GTR1 | GTR1 - Treatment interlocks VTC |  | 54752 | Y | 99629
|  |  | GTR1 | GTR1 - GTR - CBCT LFoV On Site Rec Tests |  | 55472 | Y | 99630
11 | 1 | rigging | GTR1 | GTR1 - CBCT LFoV GTR X-ray variable collimator calibration | cbct | 58334 | Y | 99631
|  |  | GTR1 | GTR1 - BDCU SD Validation test protocol (P+ C230) |  | 62555 | Y | 99632
|  |  | GTR1 | GTR1 - Abnormal beam presence C230 |  | 63748 | Y | 99633
6 | 1 | rigging | GTR1 | GTR1 - FoVCU TPO | cbct | 57211 |  | 99634
|  |  | GTR1 | GTR1 - PBS_validation_TPO-Range | VT | 18825 | Y | 99635
|  |  | GTR1 | GTR1 - TPO_PBS_Valid_Beam_Characteristics | VT | 18829 | Y | 99636
|  |  | GTR1 | GTR1 - PBSV2_BMS_BEAM_PERF_04_treatment_Time | VT | 18835 |  | 99637
|  |  | GTR1 | GTR1 - PBS_BMS_BEAM_PERF_05_Field_size.doc | VT | 18839 | Y | 99638
|  |  | GTR1 | GTR1 - PBSV2_BMS_BEAM_PERF_07_Integrated_fluence | VT | 18841 | Y | 99639
|  |  | GTR1 | GTR1 - PBS_validation_TPO-Field_conformity | VT | 18849 | Y | 99640
|  |  | GTR1 | GTR1 - PTS (renovated) operation scenario | VT | 21867 | Y | 99641
|  |  | GTR1 | GTR1 - UBTI_V1B_VT | VT | 27508 | Y | 99642
21 | 8 | validation | GTR1 | GTR1 - RPPS Increased Treatment Volume VTC | test | 34072 | Y | 99643
23 | 2 | tss | GTR1 | GTR1 - PTS RPD VTC | test | 36312 | Y | 99644
|  |  | GTR1 | GTR1 - PBS Dedicated Snout V2 VTC | VT | 39805 | Y | 99645
23 | 5 | tss | GTR1 | GTR1 - PTS CBCT VTC | test | 40374 | Y | 99646
19 | 8 | validation_pps | GTR1 | GTR1 - PTS RPD-V3 VTC | test | 55339 |  | 99647
|  |  | GTR1 | GTR1 - IEC60601-2-64 Site Tests Grade B/C: Nominal Irradiations | VT | 61260 |  | 99648
|  |  | GTR1 | GTR1 - IEC 60601-2-64 Site Test Range & Beam Line interlocks P-PLUS | VT | 61281 |  | 99649
|  |  | GTR1 | GTR1 - IEC60601-2-64 Site Tests Grade B/C: PESS | VT | 63280 |  | 99650
|  |  | GTR1 | GTR1 - Site Test Dose Monitoring Systems 60601-2-64 ProteusPLUS | VT | 66784 |  | 99651
|  |  | GTR1 | GTR1 - TC VTC OIS I/F EMRC (renovated) | VT | 26332 | Y | 99652
|  |  | GTR1 | GTR1 - Offline Review with MOSAIQ Site TPO | IT | 31095 | Y | 99653
|  |  | GTR2 | GTR2 - SRCUC cabinet signals test | UT | 57390 | Y | 99654
|  |  | GTR2 | GTR2 - SRCU Configuration Site Check integration Test | VT | 61545 | Y | 99655 | Shuling Peng
|  |  | GTR2 | GTR2 - P+ TSS validation procedure for Treatment Room PlatformD&E | VT | 61792 | Y | 99656
|  |  | GTR2 | GTR2 - Telescope (straight line) Manual Beam Tuning | VT | 7204 |  | 99657
|  |  | GTR2 | GTR2 - Preparing ESBTS-Convalgo | VT | 15288 |  | 99658
|  |  | GTR2 | GTR2 - TPO for ESS efficiency characterization | VT | 20206 |  | 99659
|  |  | GTR2 | GTR2 - PBS beam tuning | VT | 21888 |  | 99660
|  |  | GTR2 | GTR2 - Calibration of IC1 tuning gains in PBS | VT | 30685 |  | 99661
|  |  | GTR2 | GTR2 - Field regulation and volumetric repainting calibration | VT | 42969 |  | 99662
9 | 8 | function_pps | GTR2 | Correction of the offset of joint A6 of PPS LEONI | test | 68811
5 | 9 | rigging | GTR2 | GTR2 - PT project: Gantry Limit Switches Box on site unit test procedure | test | 1561 |  |  |  |  | 5% | 2
12 | 11 | function | GTR2 | GTR2 - Gantry on site integration test procedure | test | 1801 | Y | 99664
18 | 11 | validation | GTR2 | GTR2 - Gantry on site validation test procedure | test | 2125 | Y | 99665
5 | 11 | rigging | GTR2 | GTR2 - Gantry Brakes Cabinets on site unit test procedure | test | 2515 |  | 99666
5 | 12 | rigging | GTR2 | GTR2 - Gantry Drivers Cabinet on site unit test procedure | test | 5809 |  | 99667
18 | 12 | validation | GTR2 | GTR2 - GRF on-site recurrent unit tests procedure | test | 19666 |  | 99668
19 | 12 | validation | GTR2 | GTR2 - GRF On-site recurrent integration tests procedure | test | 19667 |  | 99670
20 | 12 | validation | GTR2 | GTR2 - GRF on-site recurrent validation tests procedure | test | 19668 | Y | 99671
15 | 12 | function | GTR2 | GTR2 - GRF TOC | TOC | 21058 | Y | 99672
9 | 12 | function | GTR2 | GTR2 - PCUIO unit test | test | 36877 |  | 99673
11 | 7 | function | GTR2 | GTR2 - Dedicated Nozzle Snout Holder (DNSH) IT test procedure | test | 37826 |  | 99674
15 | 7 | validation | GTR2 | GTR2 - DNSH alignment procedure using PPVS | tet | 43837 |  | 99675
4 | 8 | pps-rigging | GTR2 | GTR2 - Leoni PPS: Site Acceptance Test | test | 57306 | Y | 99676
7 | 8 | pps-rigging | GTR2 | GTR2 - Leoni PPS: Accuracy Validation | test | 57307 | Y | 99677
12 | 8 | function | GTR2 | GTR2 - Leoni Robot PPS Robot recurrent test | test | 58339 | Y | 99678
7 | 9 | rigging | GTR2 | GTR2 - Gantry balancing procedure | test | 62212 | Y | 99679 | Jianlong Wen |  | 20% | 1
25 | 7 |  | GTR2 | GTR2 - IEC60601-2-64_SiteTest_PosDevice I&D | test | 64008 |  | 99680
7 | 10 | rigging | GTR2 | GTR2 - Gantry 360 preliminary mechanical test procedure | test | 65492 |  | 99681
10 | 8 | function_pps | GTR2 | GTR2 - Leoni PPS - Update of Couch Definition File | test | 66850 |  | 99682
11 | 8 | function_pps | GTR2 | GTR2 - Leoni Robot PPS couch definition file verification | test | 64474 |  | 99683
22 | 8 |  | GTR2 | GTR2 - PMS SRIT | test | 64181 |  | 99684
4 | 5 | rigging | GTR2 | GTR2 - Flat panels GTR Mechanical structure on site unit test procedure | test | 2017 |  | 99685
6 | 5 | rigging | GTR2 | GTR2 - Flat Panels Motion Electronic Unit GTR on site unit test procedure | test | 2019 |  | 99686
16 | 6 | rigging | GTR2 | GTR2 - Nozzle X Ray tube stray radiation measurement | toc | 8218 | Y | 99687
5 | 3 | rigging | GTR2 | GTR2 - Laser Protection RVT | test | 26775 |  | 99688
18 | 6 | validation | GTR2 | GTR2 - Test of the radiographic PPVS image quality | test | 27834 | Y | 99689
18 | 1 | validation | GTR2 | GTR2 - CBCT Validation procedure | cbct | 32137 | Y | 99690
5 | 1 | rigging | GTR2 | GTR2 - Configuration procedure for the Rotoscan RS4-4 | cbct | 32138 | Y | 99691
22 | 3 |  | GTR2 | GTR2 - ProteusPlus PPVS SW: non-regression at installation | test | 67728 | Y | 99692
17 | 3 | rigging | GTR2 | GTR2 - Calibration dose vs radiographic tube settings | cbct | 34959 |  | 99693
14 | 4 | rigging | GTR2 | GTR2 - Validation of adaPT insight Geometric Calibration | test | 36048 |  | 99694
4 | 3 | rigging | GTR2 | GTR2 - 38541 CBCT IVV Radioscopic Devices Geometrical Check | test | 38541 |  | 99695
|  |  | GTR2 | GTR2 - CBCT: integration test for SRCU – PCU – X-ray Generator | IT | 43710 | Y | 99696
10 | 4 | rigging | GTR2 | GTR2 - EMD generator calibration procedure | test | 54462 |  | 99697
19 | 6 | validation | GTR2 | GTR2 - PPVS software validation - First installation | test | 54534 | Y | 99698
13 | 5 | rigging | GTR2 | GTR2 - Fine Alignment and FP calibration | test | 55357 | Y | 99699
10 | 3 | rigging | GTR2 | GTR2 - Validation of HW x-ray equipment using EMD generators | test | 56498 |  | 99700
5 | 4 | rigging | GTR2 | GTR2 - PPVS EMD installation guidelines | test | 64208 |  | 99701
4 | 4 | rigging | GTR2 | GTR2 - PPVS Hardware Inspection | test | 60180 |  | 99702
15 | 8 | validation_pps | GTR2 | GTR2 - Treatment couch recurrent validation | test | 65546 |  | 99703
|  |  | GTR2 | GTR2 - IEC60601-2-68 PMS Site tests | VT | 69408 |  | 99704
|  |  | GTR2 | GTR2 - IEC60601-2-68-Site-Data | VT | 67876 |  | 99705
|  |  | GTR2 | GTR2 - IEC60601-2-68-Site-GUI | VT | 68250 |  | 99706
|  |  | GTR2 | GTR2 - IEC60601-2-68-Site-Safety | VT | 68251 |  | 99707
|  |  | GTR2 | GTR2 - adaPT insight safety labels | VT | 68891 |  | 99708
14 | 6 | rigging | GTR2 | GTR2 - Check image orientation in AI | test | 64957 |  | 99709
6 | 3 | rigging | GTR2 | GTR2 - Testing/Validation Installation of adaPT insight | test | 66442 |  | 99710
16 | 8 | validation_pps | GTR2 | GTR2 - PPS Real Iso Offset computation | test | 32781 |  | 99711
14 | 5 | rigging | GTR2 | GTR2 - Adaptinsight_geometric_calibration_procedure | test | 36050 |  | 99712
22 | 4 |  | GTR2 | GTR2 - Clinical Installation of adaPT insight | test | 66385 |  | 99713
16 | 3 | rigging | GTR2 | GTR2 - Acquisitions for CBCT calibration | cbct | 66505 |  | 99714
16 | 2 | rigging | GTR2 | GTR2 - Creation of calibration maps for CBCT | cbct | 66519 |  | 99715
16 | 1 | rigging | GTR2 | GTR2 - Calibration of CBCT reconstruction parameters | cbct | 66530 |  | 99716
|  |  | GTR2 | GTR2 - Site integration test LPIC1 | IT | 43133 |  | 99717
24 | 7 |  | GTR2 | GTR2 - IEC60601-2-64 Site Tests Grade B/C: Accessories | test | 61271 |  | 99718
5 | 10 | rigging | GTR2 | GTR2 - GTR Nozzle Alignment Procedure | test | 15312 |  | 99719
|  |  | GTR2 | GTR2 - BDSIOS PBS Site Integration Tests | IT | 19058 |  | 99720
|  |  | GTR2 | GTR2 - SMPS Integration Test - Dekimo | IT | 45642 |  | 99721
|  |  | GTR2 | GTR2 - PBS system feedback delay calibration |  | 19376 |  | 99722
|  |  | GTR2 | GTR2 - Alignment of PBS unscanned beam |  | 21877 |  | 99723
|  |  | GTR2 | GTR2 - SAD in PBS mode |  | 21900 |  | 99724
|  |  | GTR2 | GTR2 - BDS scanning controller/scanning magnet calibration |  | 21910 |  | 99725
|  |  | GTR2 | GTR2 - PBS ICs parameters calibration |  | 22014 |  | 99726
|  |  | GTR2 | GTR2 - PBS Nozzle WET |  | 22019 |  | 99727
|  |  | GTR2 | GTR2 - IC beam size |  | 22020 |  | 99728
|  |  | GTR2 | GTR2 - PBS thresholds and safety check calibration |  | 22086 | Y | 99729
|  |  | GTR2 | GTR2 - PBS - Wiring check | WI | 33114 |  | 99730
|  |  | GTR2 | GTR2 - Fine tune distance from IC to isocenter |  | 42835 |  | 99731
22 | 2 |  | GTR2 | GTR2 - Xray Proton Colinearity for PPVS testing | test | 53920 | Y | 99732
|  |  | xxGTR2 | GTR2 - Validation procedure: target positioning in the gantry room | VT | 6712 | Y | 99733 |  | Removed
|  |  | GTR2 | GTR2 - UBTI_service_IT | IT | 27719 | Y | 99734
24 | 6 |  | GTR2 | GTR2 - Move PPDs into SETUP and TREATMENT | test | 1918 |  | 99735
23 | 11 | tss | GTR2 | GTR2 - Enable switch: emergency stop switches the hand pendant ON | test | 1929 | Y | 99736
23 | 9 | tss | GTR2 | GTR2 - Collision detection | test | 1930 | Y | 99737
23 | 8 | tss | GTR2 | GTR2 - PPD speed control Validation Test Procedure | test | 1931 | Y | 99738
23 | 7 | tss | GTR2 | GTR2 - PROXIMITY validation Tests | test | 1932 | Y | 99739
23 | 6 | tss | GTR2 | GTR2 - PPDs check during treatment | test | 1975 | Y | 99740
|  |  | GTR2 | GTR2 - BMS - PBS - Test Case 1 - TPO |  | 22058 | Y | 99741
|  |  | GTR2 | GTR2 - BMS PBS DCEU Tests |  | 22069 | Y | 99742
|  |  | GTR2 | GTR2 - Safety Check for PBS Scatterer Position |  | 36485 | Y | 99743
|  |  | GTR2 | GTR2 - Beam veto ITC | IT | 41020 | Y | 99744
|  |  | GTR2 | GTR2 - Treatment interlocks VTC |  | 54752 | Y | 99745
|  |  | GTR2 | GTR2 - GTR - CBCT LFoV On Site Rec Tests |  | 55472 | Y | 99746
11 | 1 | rigging | GTR2 | GTR2 - CBCT LFoV GTR X-ray variable collimator calibration | cbct | 58334 | Y | 99747
|  |  | GTR2 | GTR2 - BDCU SD Validation test protocol (P+ C230) |  | 62555 | Y | 99748
|  |  | GTR2 | GTR2 - Abnormal beam presence C230 |  | 63748 | Y | 99749
6 | 1 | rigging | GTR2 | GTR2 - FoVCU TPO | cbct | 57211 |  | 99750
|  |  | GTR2 | GTR2 - PBS_validation_TPO-Range | VT | 18825 | Y | 99751
|  |  | GTR2 | GTR2 - TPO_PBS_Valid_Beam_Characteristics | VT | 18829 | Y | 99752
|  |  | GTR2 | GTR2 - PBSV2_BMS_BEAM_PERF_04_treatment_Time | VT | 18835 |  | 99753
|  |  | GTR2 | GTR2 - PBS_BMS_BEAM_PERF_05_Field_size.doc | VT | 18839 | Y | 99754
|  |  | GTR2 | GTR2 - PBSV2_BMS_BEAM_PERF_07_Integrated_fluence | VT | 18841 | Y | 99755
|  |  | GTR2 | GTR2 - PBS_validation_TPO-Field_conformity | VT | 18849 | Y | 99756
|  |  | GTR2 | GTR2 - PTS (renovated) operation scenario | VT | 21867 | Y | 99757
|  |  | GTR2 | GTR2 - UBTI_V1B_VT | VT | 27508 | Y | 99758
21 | 8 | validation | GTR2 | GTR2 - RPPS Increased Treatment Volume VTC | test | 34072 | Y | 99759
23 | 2 | tss | GTR2 | GTR2 - PTS RPD VTC | test | 36312 | Y | 99760
|  |  | GTR2 | GTR2 - PBS Dedicated Snout V2 VTC | VT | 39805 | Y | 99761
23 | 5 | tss | GTR2 | GTR2 - PTS CBCT VTC | test | 40374 | Y | 99762
19 | 8 | validation_pps | GTR2 | GTR2 - PTS RPD-V3 VTC | test | 55339 |  | 99763
|  |  | GTR2 | GTR2 - IEC60601-2-64 Site Tests Grade B/C: Nominal Irradiations | VT | 61260 |  | 99764
|  |  | GTR2 | GTR2 - IEC 60601-2-64 Site Test Range & Beam Line interlocks P-PLUS | VT | 61281 |  | 99765
|  |  | GTR2 | GTR2 - IEC60601-2-64 Site Tests Grade B/C: PESS | VT | 63280 |  | 99766
|  |  | GTR2 | GTR2 - Site Test Dose Monitoring Systems 60601-2-64 ProteusPLUS | VT | 66784 |  | 99767
|  |  | GTR2 | GTR2 - TC VTC OIS I/F EMRC (renovated) | VT | 26332 | Y | 99768
|  |  | GTR2 | GTR2 - Offline Review with MOSAIQ Site TPO | IT | 31095 | Y | 99769
|  |  | GTR3 | GTR3 - SRCUC cabinet signals test | UT | 57390 | Y | 99770 | Shuling Peng |  | 90% | 64
|  |  | GTR3 | GTR3 - SRCU Configuration Site Check integration Test | VT | 61545 | Y | 99771 | Shuling Peng | QA | 100% | 4
|  |  | GTR3 | GTR3 - P+ TSS validation procedure for Treatment Room PlatformD&E | VT | 61792 | Y | 99772 | Shuling Peng
|  |  | GTR3 | GTR3 - Telescope (straight line) Manual Beam Tuning | VT | 7204 |  | 99773 | Chenchen Yuan | QA | 100% | 16
|  |  | GTR3 | GTR3 - Preparing ESBTS-Convalgo | VT | 15288 |  | 99774 | Chenchen Yuan |  | 80% | 4
|  |  | GTR3 | GTR3 - TPO for ESS efficiency characterization | VT | 20206 |  | 99775 | Chenchen Yuan
|  |  | GTR3 | GTR3 - PBS beam tuning | VT | 21888 |  | 99776 | Chenchen Yuan
|  |  | GTR3 | GTR3 - Calibration of IC1 tuning gains in PBS | VT | 30685 |  | 99777 | Chenchen Yuan
|  |  | GTR3 | GTR3 - Field regulation and volumetric repainting calibration | VT | 42969 |  | 99778 | Chenchen Yuan
9 | 8 | function_pps | GTR3 | Correction of the offset of joint A6 of PPS LEONI | test | 68811 |  |  | Deyou Wang
5 | 9 | rigging | GTR3 | GTR3 - PT project: Gantry Limit Switches Box on site unit test procedure | test | 1561 |  | 99779 | Jianlong Wen |  | 95% | 6
12 | 11 | function | GTR3 | GTR3 - Gantry on site integration test procedure | test | 1801 | Y | 99780 | Jianlong Wen |  | 35% | 8
18 | 11 | validation | GTR3 | GTR3 - Gantry on site validation test procedure | test | 2125 | Y | 99781
5 | 11 | rigging | GTR3 | GTR3 - Gantry Brakes Cabinets on site unit test procedure | test | 2515 |  | 99782 | Jianlong Wen |  | 90% | 4
5 | 12 | rigging | GTR3 | GTR3 - Gantry Drivers Cabinet on site unit test procedure | test | 5809 |  | 99783 | Jianlong Wen |  | 80% | 7
18 | 12 | validation | GTR3 | GTR3 - GRF on-site recurrent unit tests procedure | test | 19666 |  | 99784 | Deyou Wang |  | 60%
19 | 12 | validation | GTR3 | GTR3 - GRF On-site recurrent integration tests procedure | test | 19667 |  | 99785 | Deyou Wang |  | 10%
20 | 12 | validation | GTR3 | GTR3 - GRF on-site recurrent validation tests procedure | test | 19668 | Y | 99786 | Deyou Wang |  | 1%
15 | 12 | function | GTR3 | GTR3 - GRF TOC | TOC | 21058 | Y | 99787 | Deyou Wang |  | 1%
9 | 12 | function | GTR3 | GTR3 - PCUIO unit test | test | 36877 |  | 99788 | Jianlong Wen |  | 88% | 8
11 | 7 | function | GTR3 | GTR3 - Dedicated Nozzle Snout Holder (DNSH) IT test procedure | test | 37826 |  | 99789 | Jianlong Wen |  | 75% | 15
15 | 7 | validation | GTR3 | GTR3 - DNSH alignment procedure using PPVS | test | 43837 |  | 99790 | Jianlong Wen
4 | 8 | pps-rigging | GTR3 | GTR3 - Leoni PPS: Site Acceptance Test | test | 57306 | Y | 99791 | Deyou Wang
7 | 8 | pps-rigging | GTR3 | GTR3 - Leoni PPS: Accuracy Validation | test | 57307 | Y | 99792 | Deyou Wang
12 | 8 | function_pps | GTR3 | GTR3 - Leoni Robot PPS Robot recurrent test | test | 58339 | Y | 99793 | Jianlong Wen |  | 90% | 16
7 | 9 | rigging | GTR3 | GTR3 - Gantry balancing procedure | test | 62212 | Y | 99794 | Jianlong Wen |  | 95% | 21
25 | 7 |  | GTR3 | GTR3 - IEC60601-2-64_SiteTest_PosDevice I&D | test | 64008 |  | 99795 | Jianlong Wen
7 | 10 | rigging | GTR3 | GTR3 - Gantry 360 preliminary mechanical test procedure | test | 65492 |  | 99796 | Jianlong Wen |  | 90% | 30
10 | 8 | function_pps | GTR3 | GTR3 - Leoni PPS - Update of Couch Definition File | test | 66850 |  | 99797 | Jianlong Wen
11 | 8 | function_pps | GTR3 | GTR3 - Leoni Robot PPS couch definition file verification | test | 64474 |  | 99798 | Jianlong Wen
22 | 8 |  | GTR3 | GTR3 - PMS SRIT | test | 64181 |  | 99799 | Jianlong Wen
4 | 5 | rigging | GTR3 | GTR3 - Flat panels GTR Mechanical structure on site unit test procedure | test | 2017 |  | 99800 | Jianlong Wen |  | 95% | 25
6 | 5 | rigging | GTR3 | GTR3 - Flat Panels Motion Electronic Unit GTR on site unit test procedure | test | 2019 |  | 99801 | Jianlong Wen |  | 95% | 32
16 | 6 | rigging | GTR3 | GTR3 - Nozzle X Ray tube stray radiation measurement | TOC | 8218 | Y | 99802
5 | 3 | rigging | GTR3 | GTR3 - Laser Protection RVT | test | 26775 |  | 99803 | Jianlong Wen |  | 90% | 9
18 | 6 | validation | GTR3 | GTR3 - Test of the radiographic PPVS image quality | test | 27834 | Y | 99804 | Deyou Wang |  | 1%
18 | 1 | validation | GTR3 | GTR3 - CBCT Validation procedure | cbct | 32137 | Y | 99805 | Deyou Wang
5 | 1 | rigging | GTR3 | GTR3 - Configuration procedure for the Rotoscan RS4-4 | cbct | 32138 | Y | 99806 | Jianlong Wen |  | 90% | 6
22 | 3 |  | GTR3 | GTR3 - ProteusPlus PPVS SW: non-regression at installation | test | 67728 | Y | 99807 | Deyou Wang |  | 1%
17 | 3 | rigging | GTR3 | GTR3 - Calibration dose vs radiographic tube settings | cbct | 34959 |  | 99808 | Deyou Wang
14 | 4 | rigging | GTR3 | GTR3 - Validation of adaPT insight Geometric Calibration | test | 36048 |  | 99809 | Deyou Wang |  | 1%
4 | 3 | rigging | GTR3 | GTR3 - 38541 CBCT IVV Radioscopic Devices Geometrical Check | test | 38541 |  | 99810 | Jianlong Wen |  | 95% | 2
|  |  | GTR3 | GTR3 - CBCT: integration test for SRCU – PCU – X-ray Generator | IT | 43710 | Y | 99811 | Deyou Wang |  | 1%
10 | 4 | rigging | GTR3 | GTR3 - EMD generator calibration procedure | test | 54462 |  | 99812 | Jianlong Wen |  | 95% | 24
19 | 6 | validation | GTR3 | GTR3 - PPVS software validation - First installation | test | 54534 | Y | 99813 | Yang Zhao
13 | 5 | rigging | GTR3 | GTR3 - Fine Alignment and FP calibration | test | 55357 | Y | 99814 | Jianlong Wen |  | 50% | 78
10 | 3 | rigging | GTR3 | GTR3 - Validation of HW x-ray equipment using EMD generators | test | 56498 |  | 99815 | Deyou Wang |  | 1%
5 | 4 | rigging | GTR3 | GTR3 - PPVS EMD installation guidelines | test | 64208 |  | 99816 | Jianlong Wen |  | 90% | 46
4 | 4 | rigging | GTR3 | GTR3 - PPVS Hardware Inspection | test | 60180 |  | 99817 | Jianlong Wen |  | 90% | 6
15 | 8 | validation_pps | GTR3 | GTR3 - Treatment couch recurrent validation | test | 65546 |  | 99818 | Deyou Wang
|  |  | GTR3 | GTR3 - IEC60601-2-68 PMS Site tests | VT | 69408 |  | 99819 | Jianlong Wen
|  |  | GTR3 | GTR3 - IEC60601-2-68-Site-Data | VT | 67876 |  | 99820 | Yang Zhao
|  |  | GTR3 | GTR3 - IEC60601-2-68-Site-GUI | VT | 68250 |  | 99821 | Yang Zhao
|  |  | GTR3 | GTR3 - IEC60601-2-68-Site-Safety | VT | 68251 |  | 99822 | Shuling Peng
|  |  | GTR3 | GTR3 - adaPT insight safety labels | VT | 68891 |  | 99823 | Yang Zhao
14 | 6 | rigging | GTR3 | GTR3 - Check image orientation in AI | test | 64957 |  | 99824 | Jianlong Wen
6 | 3 | rigging | GTR3 | GTR3 - Testing/Validation Installation of adaPT insight | test | 66442 |  | 99825 | Yang Zhao |  | 50% | 14
16 | 8 | validation_pps | GTR3 | GTR3 - PPS Real Iso Offset computation | test | 32781 |  | 99826 | Jianlong Wen
14 | 5 | rigging | GTR3 | GTR3 - Adaptinsight_geometric_calibration_procedure | test | 36050 |  | 99827 | Deyou Wang |  | 1%
22 | 4 |  | GTR3 | GTR3 - Clinical Installation of adaPT insight | test | 66385 |  | 99828 | Yang Zhao
16 | 3 | rigging | GTR3 | GTR3 - Acquisitions for CBCT calibration | cbct | 66505 |  | 99829 | Deyou Wang |  | 1%
16 | 2 | rigging | GTR3 | GTR3 - Creation of calibration maps for CBCT | cbct | 66519 |  | 99830 | Yang Zhao
16 | 1 | rigging | GTR3 | GTR3 - Calibration of CBCT reconstruction parameters | cbct | 66530 |  | 99831 | Jianlong Wen
|  |  | GTR3 | GTR3 - Site integration test LPIC1 | IT | 43133 |  | 99832 | Yeehim Him | QA | 100% | 18
24 | 7 |  | GTR3 | GTR3 - IEC60601-2-64 Site Tests Grade B/C: Accessories | test | 61271 |  | 99833 | Yeehim Him
5 | 10 | rigging | GTR3 | GTR3 - GTR Nozzle Alignment Procedure | test | 15312 |  | 99834 | Chenchen Yuan
|  |  | GTR3 | GTR3 - BDSIOS PBS Site Integration Tests | IT | 19058 |  | 99835 | Chenchen Yuan
|  |  | GTR3 | GTR3 - SMPS Integration Test - Dekimo | IT | 45642 |  | 99836 | Chenchen Yuan |  | 50% | 48
|  |  | GTR3 | GTR3 - PBS system feedback delay calibration |  | 19376 |  | 99837 | Chenchen Yuan
|  |  | GTR3 | GTR3 - Alignment of PBS unscanned beam |  | 21877 |  | 99838 | Chenchen Yuan
|  |  | GTR3 | GTR3 - SAD in PBS mode |  | 21900 |  | 99839 | Chenchen Yuan
|  |  | GTR3 | GTR3 - BDS scanning controller/scanning magnet calibration |  | 21910 |  | 99840 | Chenchen Yuan
|  |  | GTR3 | GTR3 - PBS ICs parameters calibration |  | 22014 |  | 99841 | Chenchen Yuan
|  |  | GTR3 | GTR3 - PBS Nozzle WET |  | 22019 |  | 99842 | Chenchen Yuan
|  |  | GTR3 | GTR3 - IC beam size |  | 22020 |  | 99843 | Chenchen Yuan
|  |  | GTR3 | GTR3 - PBS thresholds and safety check calibration |  | 22086 | Y | 99844 | Chenchen Yuan
|  |  | GTR3 | GTR3 - PBS - Wiring check | WI | 33114 |  | 99845 | Chenchen Yuan
|  |  | GTR3 | GTR3 - Fine tune distance from IC to isocenter |  | 42835 |  | 99846 | Chenchen Yuan
22 | 2 |  | GTR3 | GTR3 - Xray Proton Colinearity for PPVS testing | test | 53920 | Y | 99847 | Chenchen Yuan
|  |  | xxGTR3 | GTR3 - Validation procedure: target positioning in the gantry room | test | 6712 | Y | 99848 |  | Removed
|  |  | GTR3 | GTR3 - UBTI_service_IT | IT | 27719 | Y | 99849 | Chenchen Yuan
24 | 6 |  | GTR3 | GTR3 - Move PPDs into SETUP and TREATMENT | test | 1918 |  | 99850 | Jianlong Wen
23 | 11 | tss | GTR3 | GTR3 - Enable switch: emergency stop switches the hand pendant ON | test | 1929 | Y | 99851 | Jianlong Wen
23 | 9 | tss | GTR3 | GTR3 - Collision detection | test | 1930 | Y | 99852 | Jianlong Wen
23 | 8 | tss | GTR3 | GTR3 - PPD speed control Validation Test Procedure | test | 1931 | Y | 99853 | Jianlong Wen
23 | 7 | tss | GTR3 | GTR3 - PROXIMITY validation Tests | test | 1932 | Y | 99854 | Jianlong Wen
23 | 6 | tss | GTR3 | GTR3 - PPDs check during treatment | test | 1975 | Y | 99855 | Chenchen Yuan
|  |  | GTR3 | GTR3 - BMS - PBS - Test Case 1 - TPO |  | 22058 | Y | 99856 | Chenchen Yuan
|  |  | GTR3 | GTR3 - BMS PBS DCEU Tests |  | 22069 | Y | 99857 | Chenchen Yuan
|  |  | GTR3 | GTR3 - Safety Check for PBS Scatterer Position |  | 36485 | Y | 99858 | Chenchen Yuan
|  |  | GTR3 | GTR3 - Beam veto ITC | IT | 41020 | Y | 99859 | Chenchen Yuan
|  |  | GTR3 | GTR3 - Treatment interlocks VTC |  | 54752 | Y | 99860 | Yang Zhao
|  |  | GTR3 | GTR3 - GTR - CBCT LFoV On Site Rec Tests |  | 55472 | Y | 99861 | Jianlong Wen
11 | 1 | rigging | GTR3 | GTR3 - CBCT LFoV GTR X-ray variable collimator calibration | cbct | 58334 | Y | 99862 | Jianlong Wen |  | 1% | 4
|  |  | GTR3 | GTR3 - BDCU SD Validation test protocol (P+ C230) |  | 62555 | Y | 99863 | Chenchen Yuan
|  |  | GTR3 | GTR3 - Abnormal beam presence C230 |  | 63748 | Y | 99864 | Chenchen Yuan
6 | 1 | rigging | GTR3 | GTR3 - FoVCU TPO | cbct | 57211 |  | 99865 | Jianlong Wen |  | 35% | 12
|  |  | GTR3 | GTR3 - PBS_validation_TPO-Range | VT | 18825 | Y | 99866 | Chenchen Yuan
|  |  | GTR3 | GTR3 - TPO_PBS_Valid_Beam_Characteristics | VT | 18829 | Y | 99867 | Chenchen Yuan
|  |  | GTR3 | GTR3 - PBSV2_BMS_BEAM_PERF_04_treatment_Time | VT | 18835 |  | 99868 | Chenchen Yuan
|  |  | GTR3 | GTR3 - PBS_BMS_BEAM_PERF_05_Field_size.doc | VT | 18839 | Y | 99869 | Chenchen Yuan
|  |  | GTR3 | GTR3 - PBSV2_BMS_BEAM_PERF_07_Integrated_fluence | VT | 18841 | Y | 99870 | Chenchen Yuan
|  |  | GTR3 | GTR3 - PBS_validation_TPO-Field_conformity | VT | 18849 | Y | 99871 | Chenchen Yuan
|  |  | GTR3 | GTR3 - PTS (renovated) operation scenario | VT | 21867 | Y | 99872 | Yang Zhao
|  |  | GTR3 | GTR3 - UBTI_V1B_VT | VT | 27508 | Y | 99873 | Yang Zhao
21 | 8 | validation | GTR3 | GTR3 - RPPS Increased Treatment Volume VTC | test | 34072 | Y | 99874 | Yang Zhao
23 | 2 | tss | GTR3 | GTR3 - PTS RPD VTC | test | 36312 | Y | 99875 | Yang Zhao
|  |  | GTR3 | GTR3 - PBS Dedicated Snout V2 VTC | VT | 39805 | Y | 99876 | Yang Zhao
23 | 5 | tss | GTR3 | GTR3 - PTS CBCT VTC | test | 40374 | Y | 99877 | Yang Zhao
19 | 8 | validation_pps | GTR3 | GTR3 - PTS RPD-V3 VTC | test | 55339 |  | 99878 | Yang Zhao
|  |  | GTR3 | GTR3 - IEC60601-2-64 Site Tests Grade B/C: Nominal Irradiations | VT | 61260 |  | 99879 | Chenchen Yuan
|  |  | GTR3 | GTR3 - IEC 60601-2-64 Site Test Range & Beam Line interlocks P-PLUS | VT | 61281 |  | 99880 | Chenchen Yuan
|  |  | GTR3 | GTR3 - IEC60601-2-64 Site Tests Grade B/C: PESS | VT | 63280 |  | 99881 | Chenchen Yuan
|  |  | GTR3 | GTR3 - Site Test Dose Monitoring Systems 60601-2-64 ProteusPLUS | VT | 66784 |  | 99882 | Chenchen Yuan
|  |  | GTR3 | GTR3 - TC VTC OIS I/F EMRC (renovated) | VT | 26332 | Y | 99883 | Yang Zhao
|  |  | GTR3 | GTR3 - Offline Review with MOSAIQ Site TPO | IT | 31095 | Y | 99884 | Yang Zhao
5 | 8 | rigging | GTR1 | Leoni PPS Accuracy Calibration | test | 57309
5 | 8 | rigging | GTR2 | Leoni PPS Accuracy Calibration | test | 57309
5 | 8 | rigging | GTR3 | Leoni PPS Accuracy Calibration | test | 57309
5 | 5 | rigging | GTR1 | PPVS Installation Follow Up | test | 55262
5 | 5 | rigging | GTR2 | PPVS Installation Follow Up | test | 55262
5 | 5 | rigging | GTR3 | PPVS Installation Follow Up | test | 55262
20 | 6 | validation | GTR1 | upgrade testing PPVS SW | test | 75953 | Y |  | Yang Zhao
20 | 6 | validation | GTR2 | upgrade testing PPVS SW | test | 75953 | Y |  | Yang Zhao
20 | 6 | validation | GTR3 | upgrade testing PPVS SW | test | 75953 | Y |  | Yang Zhao

## Comments
Row 1 | INSTRUCTIONS: Be detailed Every time you do work related to a test procedure, you will record your comments, hours used, hours lost, and reasons for time lost on this sheet. Remember that Hours Used is total duration, and Hours Lost is just a part of that duration. So you should never record Hours Lost without recording Hours Used. | BJones 1-USA | 01/26/17 5:12 AM
Row 3 | waiting for Gefos send taking over certificate. | Sherman Chang | 07/22/20 10:07 AM
Row 4 | waiting for Gefos send taking over certificate. and summary report | Sherman Chang | 07/22/20 10:07 AM
Row 5 | Bug assigned to @haiyuan.pan@iba-group.com , please run the procedure and update in STM. | Sherman Chang | 07/22/20 10:07 AM
Row 6 | Bug assigned to @haiyuan.pan@iba-group.com , please run the procedure and update in STM. | Sherman Chang | 07/22/20 10:08 AM
Row 7 | Bug assigned to @haiyuan.pan@iba-group.com , please run the procedure and update in STM. | Sherman Chang | 07/22/20 10:08 AM
Row 10 | The WCR rigging is almost done, we just have to finalise the hose connections between the building and the conditionner. | Pierre Collin | 08/14/19 4:55 PM
Row 11 | All the magnet are aligned and today Karel will align the accessories | Pierre Collin | 09/05/19 12:03 PM
Row 12 | to complete this procedure, the customer have to put the grouting. | Pierre Collin | 10/01/19 9:14 AM
Row 12 | All the tables are install, still the grout the tables | Pierre Collin | 08/14/19 4:57 PM
Row 13 | The installation and alignment are 100% done | Pierre Collin | 08/14/19 5:07 PM
Row 14 | run water through Gantries, in room 3 we have to connect pipes inside two magnets. also need purchase new deionization bottles | Andriy Lytovchenko | 11/22/19 10:14 PM
Row 14 | Cables for pumps arrived, pulled and connected. cables for sensors ongoing. Problem on customer chiller side. pressure is not enough for cool down our system. flushing and cleaning chiller side. | Andriy Lytovchenko | 09/19/19 4:50 PM
Row 14 | Pressure test run with several leakage found under gantry, beam line. fixed all leakages. pressure test with 12 bar passed. Water tank filled. Pumps with motors aligned. waiting cables connection with pumps for start flashing. | Andriy Lytovchenko | 08/26/19 5:13 PM
Row 14 | the start to weld the pipes to connect the Cyclo water manifold | Pierre Collin | 08/14/19 5:14 PM
Row 14 | We order the stainless steel, we don't want to use the PVC one | Pierre Collin | 08/14/19 5:14 PM
Row 15 | M.C. P.S is water connected | Pierre Collin | 10/01/19 9:16 AM
Row 15 | the flush done and the pipe are connected to the M.C. | Pierre Collin | 10/01/19 9:16 AM
Row 15 | flushing done. Problem on customer chiller side. pressure is not enough for cool down our system. customer is flushing and cleaning chiller pipes. | Andriy Lytovchenko | 09/19/19 4:52 PM
Row 15 | customer valves exchanged for correct max pressure 25 bar. pressure test passed. pipes on cyclotron side connected ready for flashing test. pumps power cables still not delivered for run flushing. | Andriy Lytovchenko | 09/02/19 9:01 PM
Row 15 | Pressure test no pass. got leakage near cyclotron. valves installed by customer had maximum pressure 16 bar. New valves for 25 bar ordered and will install soon than test will run again. | Andriy Lytovchenko | 08/26/19 5:07 PM
Row 15 | the riggers start to weld the cooling pipe | Pierre Collin | 08/14/19 5:10 PM
Row 15 | we order the pipe,. | Pierre Collin | 08/14/19 5:09 PM
Row 16 | done | Pierre Collin | 10/01/19 9:17 AM
Row 16 | Flashing finished. filters cleaned. ODP cooling water manifold connected. ready for use after all cooling groups finished and customer run Chiller. | Andriy Lytovchenko | 09/02/19 8:56 PM
Row 16 | run pressure test. all pipes keep 6 bar 16 hours no leakage found. water tank filled and flashing ongoing. | Andriy Lytovchenko | 08/26/19 4:59 PM
Row 17 | done | Pierre Collin | 10/01/19 9:18 AM
Row 18 | done | Pierre Collin | 10/01/19 9:18 AM
Row 18 | Pressure test done and flush done | Pierre Collin | 09/05/19 12:05 PM
Row 19 | Done at 21H00 | Pierre Collin | 08/14/19 5:15 PM
Row 21 | @yang.zhao@iba-group.com , Please follow up with customer about the network cabling certificates. | Sherman Chang | 07/22/20 2:58 PM
Row 21 | Still waiting for the certifications. | Sherman Chang | 06/22/20 10:11 AM
Row 21 | Communicated with customer, asking for the certificates. Waiting feedback. | Sherman Chang | 04/20/20 5:11 PM
Row 22 | the cyclo alignment is 100% done | Pierre Collin | 08/14/19 5:21 PM
Row 23 | 2020-7-22 With Haiyuan, cable correction for MCR | Shuling Peng | 07/22/20 6:54 PM
Row 23 | 2020-7-21 With Haiyuan, cable correction for WCR | Shuling Peng | 07/21/20 5:53 PM
Row 23 | 2020-7-20 With Haiyuan, cable correction for PSR | Shuling Peng | 07/21/20 5:52 PM
Row 23 | 2020-7-14 Work with Mathieu remotely & Haiyuan, cable inspection for PSR & WCR & MCR & GTR3 & TCR3 | Shuling Peng | 07/15/20 6:08 PM
Row 23 | 2020-7-8 Work preparation for MID 51093 & MID 7188 & MID 7189 | Shuling Peng | 07/08/20 5:34 PM
Row 23 | Add bridge between XU.2.04.1 – XU.2.04.2 to get power for MCR Cabinet 1 regarding BUG 105034. Removed the cable WB241 from QU2.75 to QU2.77 in the PDC cabinet +G01 Pulled a new cable WB240 with a section of 3G2,5 from QU2,75 +G01 to +B11 – X0. | Andriy Lytovchenko | 12/10/19 6:21 PM
Row 23 | We found that we can’t ON Power for MCR Cabinet 1 inside PDC Cyclo Cabinet, It has to ON by relay KU2.04 which manage through KT00 by signal Emergency Stop TCR (XU2.04.2). We temporary plugged it to XAU 6 (Emergency Stop STD+EMG) which comes from SRCU0. Waiting response for correction from LLN. | Andriy Lytovchenko | 12/09/19 5:19 PM
Row 23 | MPSS switching cabinet hadn't connect to PDC's, no wiring list for this. | Andriy Lytovchenko | 12/09/19 5:16 PM
Row 23 | Work outside our scope (not improved yet for Guangzhou site) : all the grounding cables for all the cabinets, Stands, cables trays (not installed by the BT), metal support,... Lost hour!? I made a plan for all the new grounding and provided to R&D. They are ongoing to improve the grounding cabeling for the next sites. | Mathieu Eppe | 10/17/19 10:25 AM
Row 23 | Hours used update until monday 14-10 | Mathieu Eppe | 10/17/19 10:12 AM
Row 23 | 15.10.2019 same remarks one month later!!! Why I come on site if all the cables aren't there? Very hard to plan the work without devices! We are very blocked because we waiting some cables from LLN. We need to pull some cables missing to continue with all the rest. We must respect the order of the cables inside the Gore Rack (energy chain) otherise the clamps can not tight correctly the cables. - Huge difficulties to receive the missing cables earlier! Customs in China. A lot of delay for the Electrical installatio | Mathieu Eppe | 10/16/19 10:07 AM
Row 23 | 20.09.19 - We are very blocked because we waiting some cables from LLN. We need to pull some cables missing to continue with all the rest. We must respect the order of the cables inside the Gore Rack (energy chain) otherise the clamps can not tight correctly the cables. - Huge difficulties to receive the missing cables earlier! Customs in China. A lot of delay for the Electrical installation. | Mathieu Eppe | 09/20/19 10:39 AM
Row 23 | 11.09.19 Arrival time sub-contractor / Departure: 8h00 - 18h No time to writhe my reports. - A lot of delay due to the cables missings. I try to have a clear status with the cables ordered locally, from LLN, from others china sites... We are blocked in some area because the cables are missing. A lot of devices in the gantry rooms are not installed, we cann't connect or pull these cables. The customer didn't pull yet the HK from the "boxes" to our cabinets. Our sub contractor needs to pull them... Not our job to do that. The sub contractor worked and work every saturday. Nobody from IBA is on site during the week end! Normal or Not? | Mathieu Eppe | 09/11/19 2:49 PM
Row 23 | 15.08.19 Arrival time MEP / Departure: 8h10 - 17h30 (+50min from the hotel to the site and the same to coming back) Arrival time sub-contractor / Departure: 8h00 - 18h - No report during a few days because too much work. - They start to do the +95 and +93 cabinet for the cyclo. - They finish with the connector inside the PDC cabinet. Some cales are missing because the cables arrive later with the UK. (No one knows when it should be delivery on site) - I lost a lot hours due to the internet!! it's horrible. I can't have access to smartsheet, outlook, sometimes internet is off for a long time... - A lot of difficulties with the cables, because the site is old, they ordered them 3 years ago. Now we use a new WL with Eplan. A lot of difference between the cables on site, the cables needed and the cables in the WL or EPLAN. - We waiting the shipment from Belgium. lost time because we need to plan the work at the different zone because some cables arn't there... - We have some difficulties with the labeling inside the cabinet. They are old and with the old labeling. | Mathieu Eppe | 08/15/19 10:50 AM
Row 23 | 08.08.19 Arrival time MEP / Departure: 8h00 - 17h30 (+50min from the hotel to the site and the same to coming back) Arrival time sub-contractor / Departure: 8h00 - 18h - They continue and finish with the pulling cables in the Gantry 2 to the Power junction box +G99 - They started to remove the scaffolding and move him in the Gantry 1. - They continues to pull some others cables. - I tried to gather all the information in the WL. - I saw a few hazardous situations. False floor not attached correctly on the floor, HK braker switch on side customer, work in height not really safely,... I didn't have time to do an EHS report yet. - Leo asked me to make a plan for more or less two work weeks and provide that to "Better". Of course the plan can change any days. - I asked to the the Site tech, facility inspector, to have a status about the HK. Because the customer installed the extra boxes with others brakers on the wall but the cables between theses boxes and our cabinets are missing. - | Mathieu Eppe | 08/08/19 12:56 PM
Row 23 | 07.08.19 Arrival time MEP / Departure: 8h00 - 17h30 Arrival time sub-contractor / Departure: 8h00 - 18h - Safety Induction - Site tour - short meeting with the sub-contractor "Beter" - I checked if they had enough work to do. - they pull the power cables from PSR to the Gantry 2 on top of the Pit. They built a scaffolding to have access on top. | Mathieu Eppe | 08/08/19 12:50 PM
Row 24 | Waiting for PCUs for GTR1/2 | Sherman Chang | 07/17/20 4:16 PM
| So any not fully finished procedure, will reset back to 99%. | Sherman Chang | 07/22/20 10:08 AM
Row 24 | due to the new PSR design, we have to ask the building to add one support frame | Pierre Collin | 09/05/19 12:07 PM
Row 24 | PSR cabinet are 100% done, still some cabinet inside the gantry (they should be install after alignment | Pierre Collin | 08/14/19 5:22 PM
Row 25 | please @shuling.peng@iba-group.com update status. | Sherman Chang | 07/17/20 4:17 PM
| The procedure is only done for GTR3. The cable connection for GTR1 & GTR2 are not done by customer yet. | Shuling Peng | 07/17/20 5:28 PM
Row 25 | May 18th,2020. 1) Rearrange cables for SRCU3 2) Under the suggestion of cotrol team, I upgraded BLCU code to the latest version, BLCU-B2.0a. | Shuling Peng | 05/18/20 10:53 PM
Row 25 | April 7th,2019. Bug 109056 is done. Besides, the relay box for SRCUi has been installed, but the power cables for those still are still to be installed | Shuling Peng | 04/08/20 8:14 AM
Row 26 | @shuling.peng@iba-group.com please update status. | Sherman Chang | 07/17/20 4:17 PM
| The TSS installation & upgrade is done in GTR3. Since the equipment for TSS is not completely done in GTR1 & GTR2, the procedure is not done for GTR1,2 yet. | Shuling Peng | 07/17/20 5:19 PM
Row 28 | I ask to add one support to help the futur rigger to rig the nozzle lifting beam | Pierre Collin | 09/05/19 12:08 PM
Row 29 | still the axial restraint alignment | Pierre Collin | 10/01/19 10:12 AM
Row 29 | Taking over protocol was signed by me on 06/09/19 | Mikhail BADYLEVICH | 09/06/19 2:37 PM
Row 29 | I design a new grounding barre | Pierre Collin | 09/05/19 12:12 PM
Row 29 | I design a new roughing pump support to prevent any futur EHS ticket and pump damages | Pierre Collin | 09/05/19 12:11 PM
Row 30 | Because the survey report just been attached yesterday. Status set back to CCB, need review during next CCB meeting. | Sherman Chang | 07/22/20 5:36 PM
Row 30 | Taking over protocol was signed by me on 06/09/19 | Mikhail BADYLEVICH | 09/06/19 2:36 PM
Row 30 | we have a problem with the flat panel fixation bolts, when we try to torque them some of those brake. I think we have a problem with the bolt quality. | Pierre Collin | 09/05/19 12:15 PM
Row 32 | We lost a LOT of times with the new packing list, we must go back to the previus one. I don't know how to do in Beijing. | Pierre Collin | 09/05/19 12:19 PM
Row 34 | It's painfull to drill and tap the PPS feet Bolts, I'll design to tools to help the riggers and try to safe a times. | Pierre Collin | 09/05/19 12:22 PM
Row 37 | @shuling, please create related bug in bugzilla | Chenchen Yuan | 07/21/20 3:39 PM
Row 37 | Add MID 72997 in ML2B. From Rodrigue Glineur: Missing procedure for recurrent test of installation of WMEBCU (HPV3). Bug 96254 | David Ringenbach | 09/28/18 7:30 PM
Row 39 | I ask to add one support to help the futur rigger to rig the nozzle lifting beam | Pierre Collin | 09/05/19 12:09 PM
Row 40 | Taking over protocol was signed by me on 06/09/19 | Mikhail BADYLEVICH | 09/06/19 2:37 PM
Row 40 | I design a new grounding barre | Pierre Collin | 09/05/19 12:13 PM
Row 40 | I design a new roughing pump support to prevent any futur EHS ticket and pump damages | Pierre Collin | 09/05/19 12:11 PM
Row 41 | Because the survey report just been attached yesterday. Status set back to CCB, need review during next CCB meeting. | Sherman Chang | 07/22/20 5:36 PM
Row 41 | Taking over protocol was signed by me on 06/09/19 | Mikhail BADYLEVICH | 09/06/19 2:35 PM
Row 41 | we have a problem with the flat panel fixation bolts, when we try to torque them some of those brake. I think we have a problem with the bolt quality. | Pierre Collin | 09/05/19 12:16 PM
Row 43 | We lost a LOT of times with the new packing list, we must go back to the previus one. I don't know how to do in Beijing. | Pierre Collin | 09/05/19 12:23 PM
Row 45 | It's painfull to drill and tap the PPS feet Bolts, I'll design to tools to help the riggers and try to safe a times. | Pierre Collin | 09/05/19 12:23 PM
Row 48 | @shuling, please create related bug in bugzilla | Chenchen Yuan | 07/21/20 3:38 PM
Row 48 | Add MID 72997 in ML2c. From Rodrigue Glineur: Missing procedure for recurrent test of installation of WMEBCU (HPV3). Bug 96254 | David Ringenbach | 09/28/18 7:32 PM
Row 50 | I ask to add one support to help the futur rigger to rig the nozzle lifting beam | Pierre Collin | 09/05/19 12:09 PM
Row 51 | Taking over protocol was signed by me on 06/09/19 | Mikhail BADYLEVICH | 09/06/19 2:38 PM
Row 51 | I design a new grounding barre | Pierre Collin | 09/05/19 12:13 PM
Row 51 | I design a new roughing pump support to prevent any futur EHS ticket and pump damages | Pierre Collin | 09/05/19 12:11 PM
Row 52 | Because the survey report just been attached yesterday. Status set back to CCB, need review during next CCB meeting. | Sherman Chang | 07/22/20 5:36 PM
Row 52 | Taking over protocol was signed by me on 06/09/19 | Mikhail BADYLEVICH | 09/06/19 2:36 PM
Row 52 | we have a problem with the flat panel fixation bolts, when we try to torque them some of those brake. I think we have a problem with the bolt quality. | Pierre Collin | 09/05/19 12:16 PM
Row 54 | We lost a LOT of times with the new packing list, we must go back to the previus one. I don't know how to do in Beijing. | Pierre Collin | 09/05/19 12:23 PM
Row 56 | Sherman .C Pressure air supply system not in BOM . Bug109028 created. We borrowed the one for DNSH in GTR2 temporary | Douglas Zhao | 09/16/19 5:00 PM
Row 56 | It's painfull to drill and tap the PPS feet Bolts, I'll design to tools to help the riggers and try to safe a times. | Pierre Collin | 09/05/19 12:24 PM
Row 57 | finish hardware installation and test by PMS; need paper work; | Chenchen Yuan | 05/31/20 2:37 PM
Row 58 | 2020-7-17 With Zhao Yang, There is a missing cable in WMEBCU. The problem was solved by installing the missing cable. | Shuling Peng | 07/17/20 5:54 PM
Row 58 | 2020-7-16 With Zhao Yang, We found a problem that when we pressed the Enable button, the RPD AP on the WMEBCU - VNC interface showed no response. Need control team suppoet. | Shuling Peng | 07/17/20 5:53 PM
Row 59 | need create related bug in bugzilla | Chenchen Yuan | 07/21/20 3:38 PM
Row 59 | MID72997 finish about 15% | Deyou Wang | 04/16/20 10:45 PM
Row 59 | MID 53256 finished about 70%. | Deyou Wang | 04/15/20 11:21 PM
Row 59 | work with Jun,1.change the PS for laser 5 and other lasers from PCU cabinet side, put rack back to cabinet. 2.put wires burndy connetor for WB3333,WB3366,WB 3462,WB3367,WB3363,ON WMEBCU cabinet side,verify them. | Deyou Wang | 03/31/20 11:34 PM
Row 59 | Add MID 72997 in ML2d. From Rodrigue Glineur: Missing procedure for recurrent test of installation of WMEBCU (HPV3). Bug 96254 | David Ringenbach | 09/28/18 7:32 PM
Row 62 | JAN, 14th. working with cable connection for SRCU0 relay box | Shuling Peng | 01/15/20 1:06 PM
Row 62 | JAN, 9th. Done with SRCU0 relay box installation | Shuling Peng | 01/14/20 8:43 AM
Row 62 | DEC,24th. With Jeff. Done with SRCU0 signal test except for PDC test | Shuling Peng | 12/25/19 9:25 AM
Row 62 | With Shuling, PDCs connected to SRCU0. correct response for Emergency Stop buttons. Remind to Check Hook UP 2 when Customer upgrade it. BAEU/BIREU working with BCP and response to TSS correctly. | Andriy Lytovchenko | 12/09/19 5:06 PM
Row 62 | with Shuling, spend time for check errors on Analog out PLC modules. attached relays straight to plc, attached resistors, swap modules with no error one, nothing help. When start PLC, no errors and errors comes after 6 minutes. | Andriy Lytovchenko | 11/22/19 10:04 PM
Row 62 | confirmed issue with customer for HK1,2. Customer going to change breakers in 2 days. 400 V for HK3,4,5 On. | Andriy Lytovchenko | 11/20/19 10:55 PM
Row 62 | with Shuling, | Andriy Lytovchenko | 11/20/19 10:52 PM
Row 62 | Tested wiring between SRCU0 and Hook ups in PSR, HK 3,4 MPSS Dipole,Qad. have feedback in serial. Also there are no 400V, need check with Better and customer. HK 6,7 PDCs have feed backs. HK.1 has shortcut with 24V ..need investme with better and customer. HK 2 looks like not correct type beaker installed which doesn't repeat to our RSS correctly. | Andriy Lytovchenko | 11/20/19 12:22 AM
Row 62 | The relays for bells and signal lights were installed last Friday. However, we found the power supply for the relays were not compatible. I requested the customer to change the power supple tomorrow. | Shuling Peng | 11/18/19 11:28 PM
Row 62 | Operational software added to memory card, SRCU connected to Archestra, with help of Control team it appeared on Scada screens. Customer will complete wiring with search buttons (one remind in volt), door switches not connected yet, doors not installed on beam line corridor yet. cabling between SRCU0 and Hook-ups of PDC not done. (need install additional box in PSR for connect cables between). | Andriy Lytovchenko | 09/26/19 1:41 PM
Row 62 | together with Kang and Mohan, loaded last software version of SRCU0 into flesh card, installed card into PLC, run wiring test and find bug in software, one of input commented wrong. waiting update from A team. buildings doors mechanical switches need exchange with magnetic switches. | Andriy Lytovchenko | 08/13/19 9:16 AM
Row 63 | Reviewed by CCB T30: xwchang, ccyuan. OK to close and set to QA for validation. | Sherman Chang | 07/22/20 10:11 AM
Row 66 | Reviewed by v&v ccb t30: xwchang, ccyuan: check the Tools information from Bug comments. OK to closed from site, set to QA for validation. | Sherman Chang | 07/23/20 10:20 AM
Row 66 | Necessary tools information missing, bug assigned to Shuge for investigation. | Sherman Chang | 03/19/20 2:02 PM
Row 66 | 19 Nov 2019, Wed, afternoon shift notes added by MWE, Pascal and Zihua (Zihua joined afternoon shift, today). "We investigated the radial probe TSS switch mismatch. We found no mismatch. The switches are working properly until the SRCU0. SRCU0 LED front face (I30.6 & I32.2) acted as expected. Thus, we think that the problem is software. We tried by installing the bypass connector to simulate the SRCU1 – 2 -3 ok feedback. The problem remains." added 1 hr used. | Mike Welch | 11/21/19 7:56 PM
Row 66 | 20 Nov 2019, Wed, on shift with Shuge. We checked that there is contact touch with the RP using the tape method. First test failed. We adjusted the Upper RF brackets from 260mm to 262mm, toward the middle plane from the hill valley (upper, in this case) for grounding the RP. We checked that the contact does now touch the RP, using the tape method when closing the Cyclo. Cleaned all surfaces and closed the Cyclo for the day. | Mike Welch | 11/21/19 6:58 PM
Row 66 | 19 Nov 2019, Tues, on shift with Shuge. Due to questions about Archestra parameters archived vs current parameters, checked and adjusted Radial Probe Init position. New Init Position is confirmed, 1265 mm. Verified measurements at, 1030, 1000, 300 and full in are valid, based on the new Init position stored, We also checked the switches function of the RP at Center, i30.6 RP Cent NC and i32.2 RP Cent NO. Our initial results show i30.6 is OK, i32.2 results were NOK during our testing, but following shift commented that these were also OK. Possibly, the time delay for reading the WEB results was not followed closely. TSS enable still in question and not resolved, but possibly not due to RP CENT NO/NC switches and timing. | Mike Welch | 11/21/19 6:56 PM
Row 66 | We adjusted the lower RF brackets grounding the RP. We adjusted also the upper one at same level (versus the valley yoke surface). We checked that the contact touch the RP with tap. It did not. | Pascal PELERIN | 11/20/19 1:56 PM
Row 66 | Cooling pipes on the radial have to be connected. | Rene BREEMANS | 09/30/19 10:47 AM
Row 66 | MID finished | Jean-Michel Denis | 09/11/19 4:39 PM
Row 66 | only the step inside the cyclo are left, everything else is done | Jean-Michel Denis | 09/10/19 5:40 PM
Row 66 | Endless screw greased, Radial probe oiled, and initialized positioned roughly calibrated at 1260mm (instead of 1270mm) | Jean-Michel Denis | 09/10/19 8:54 AM
Row 67 | Reviewed by CCB T30: xwchang, ccyuan. OK to close and set to QA for validation. | Sherman Chang | 07/22/20 10:14 AM
Row 67 | Where's the procedure? | Sherman Chang | 06/29/20 9:27 AM
Row 67 | work with Shuge and Jinan: 1. With the help of Haiyuan, connect dry air to IC cyclo start flushing 2. Rebuild and upgrade Source#3 | Lipeng Han | 04/29/20 5:57 PM
Row 67 | Rebuild Source#1 and upgrade it according to BUG 88089, fix the water leak. Source #1 efficiency is higher than source#2 | Lipeng Han | 04/28/20 3:55 PM
Row 67 | we leak check the source (source n°3). We found 1 leak at the stack of hydrogen fittings (on external hydrogen line). We found maybe a leak at the peek fitting (on source peek block at the shaft basement). | Pascal PELERIN | 11/18/19 12:09 PM
Row 67 | New leak detection done: no leak. Therefore, the leak is at the dee maybe. Try to build an homemade DN40 to 10/12 pipe adaptor. It did not work (welding of 2 different materials not possible by local welders). Chenchen found an adaptor on internet. He purchased 3 items. | Pascal PELERIN | 10/30/19 9:21 AM
Row 67 | New leak detection done at H3 upper pillars: still leak. Open cyclo, remove dee, set 2 caps at the cooling pillars / dee side. Close cyclo | Pascal PELERIN | 10/30/19 9:19 AM
Row 67 | H3 upper dee has been dismantled. O-ring seals have been cleaned. Reassembly with new O-rings & new RF springs. Close cyclo for pumping. | Pascal PELERIN | 10/30/19 9:18 AM
Row 67 | We found 2 leaks. 1 leak at the source VAT valve, 1 leak at cavity pillar - upper - H3 side. We need to open cyclo for fixing this leak. | Pascal PELERIN | 10/17/19 11:44 AM
Row 67 | After the troubleshooting the source movement issues, continued the test procedure. TPG adjusted for the source are accurate now. Finished the TPO dt 3rd OCT 2019 | Mohan Naikare | 10/04/19 1:11 PM
Row 67 | After the troubleshooting the source movement issues, continued the test procedure. TPG adjusted for the source are accurate now. Finished the TPO dt 3rd OCT 2019 | Mohan Naikare | 10/04/19 1:11 PM
Row 67 | There were two issues observed in the source movement. One was the ACU code was not OK to move the source from the middle limit switch because the VAT valve was always in closed mode. This was corrected by the control team. Later the azimutal motor movement was not correct because the upper limit switch was not closing. Adjust the s/w and now source movement IN/OUT is good. Dt : 2nd OCT 2019 -2nd half | Mohan Naikare | 10/04/19 9:28 AM
Row 67 | Problem faced in source movement. After reseting the power several time & connection check homing of source woke. Source Vacuum TPG set for UR. But after the OUT command the failure happen in source movement. Dt : 30th Sept & 1st OCT 2019. 12 noon | Mohan Naikare | 10/01/19 1:47 PM
Row 68 | Reviewed by CCB T30: xwchang, ccyuan. OK to close and set to QA for validation. | Sherman Chang | 07/22/20 10:14 AM
Row 68 | Where's the procedure | Sherman Chang | 06/29/20 9:27 AM
Row 68 | reoptimize beam at extraction due to deflector exit position out of calbration | Lipeng Han | 06/15/20 9:05 PM
Row 68 | 1. Change deflector minimum decrease limit from 0.12 to 0.2mA. 2. Found radial probe could not move, restart beckoff driver twice then OK, then the BCM1C show device status error and modbus error, power cyclo doesn`t work, have to restart ACU, that will take more than half hour to recover system. 3. Try to shoot beam after a whole night condition of deflector, still the same result: deflector start to arc as soon as we shoot beam. 4. Since we could not open cyclo, we restart deep condition the deflector, huge are during two hours, will leave the deflector condition another night. 5. Found the YLCU ethernet cable not connected, try to connect it to the port on the wall, doesn’t work, ask get help from Bruce, we change the port from tu29 to tu31, then it`s OK 6. ACU cabinet roof fan broken, need to fix it. Ask Chenchen if we have spare parts, seems not. | Lipeng Han | 04/27/20 5:20 PM
Row 68 | Try to shoot beam on BS, deflector start to arc, decide to deep condition it. Set to 54KV and oxygen fully open and keep it whole night. Before we starting condition, Shuling came to downgrade the TSS to original version, we stop everything and restart, need to restart the whole PDC. took about one hour to recover it. | Lipeng Han | 04/27/20 8:32 AM
Row 68 | 8h chart has been annexed to the procedure. Procedure signed & put on the Mikhail desk. | Pascal PELERIN | 11/28/19 5:51 PM
Row 68 | deflector conditioning has been started on 25th of November 2019. No problem, the system reached 50kV without any issue. We played the procedure 1760. Need the 8h chart before closing the procedure. | Pascal PELERIN | 11/26/19 5:11 PM
Row 68 | 21 Nov 2019, Thurs, on shift with Shuge. Filled Deflector Resistor Oil Box with, 76.03.02.0029 oil. | Mike Welch | 11/21/19 7:46 PM
Row 68 | 19 Nov 2019, Tues, on shift with Shuge. Due to questions about archived vs current parameters data, confirmed/verified Deflector measurements at position indicated to be true. Archestra entrance/exit=1052.5/1067.0 mm. Using Caliper on a stick, 1052.78/1068.6 mm. IBA Stick, 33.72/45.5 mm, measurements collected. Gaps 2.5/2.5 mm. | Mike Welch | 11/21/19 6:20 PM
Row 68 | fulfill the water line. We found 1 leak at the fitting / entrance side. We tight the fitting: no more leak after that. | Pascal PELERIN | 10/30/19 5:03 PM
Row 68 | Add N connector at PS side - HV cable. Add N connector at oil box side - HC cable. Screw 400k resistor to the HV cables of the oil box. Fulfilled box by oil. | Pascal PELERIN | 10/10/19 5:10 PM
Row 68 | add 1 connector (61.03.03.0153) at the HV line, deflector side. | Pascal PELERIN | 10/09/19 5:05 PM
Row 68 | Finished the Deflector movement calibration. All mechanical test for MID1760 are done | Jean-Michel Denis | 09/17/19 5:05 PM
Row 68 | Deflector was cleaned, the blade insulator were placed with Indium sheets on top and bottom of both insulators. Gap set to 2.45mm at the entrance and 2.45mm at the exit. Limit switches position of the deflector movement were adjusted. Entrance movement was calibrated. | Jean-Michel Denis | 09/17/19 8:56 AM
Row 68 | Deflector movement shaft have been place, HV cable have been redone, good electrical contact between blade and cable now. Carrot vacuum tested, no leak | Jean-Michel Denis | 09/10/19 9:17 AM
Row 69 | Work half day on troubleshooting profibus issue, finally find out the yoke heater cabinet profibus cable was switch to on, which should be off.Now all cooling can turn on, but general cooling pressure is not enough, find out the bypass on beam line cooling is on, switch off the bypass valve.pressure reach 8 bar | Lipeng Han | 10/22/19 9:49 PM
Row 69 | Section 7.4 -Pg 1 MC cooling security. Have a cosmetic change it show RED for the MCPS water flow symbol which is to be yellow. Reported this to control team and they will make the necessary change. | Mohan Naikare | 10/07/19 4:04 PM
Row 69 | Finished the Water cooling integration test. IN section 7.3 of C230-cooling-IVVP-175.3 --Pg 11. Only pressure output too high alarm can not be simulated because we not allowed to use value below 5 bar pressure. And its not recommended to run the pump above 10 bar pressure just for observing the high pressure alarm. | Mohan Naikare | 10/07/19 3:57 PM
Row 70 | Wait for customer buying new ball valve to change, but still not get any reply. next, I'm plan to sent the flow meter to factory for calibration and clean heat exchanger. | Haiyuan Pan | 07/20/20 6:07 PM
Row 70 | @haiyuan.pan@iba-group.com Please update status | Sherman Chang | 07/20/20 9:48 AM
Row 70 | Lost a shift because the building water being really dirty, clogged the strainer before the exchangers. The customer not being available, we had to flush their circuit ourselves | Jean-Michel Denis | 01/23/20 12:06 PM
Row 70 | 01.16.2020. with Heyuan. Cleaned main strain of General conditioner. GTR1,2,3 strains. Corrected drivers settings. adjusted Relief valve V194 for 2.8 Bar. Even we run water trough all rooms and BTL with bypass of beamline closed we still have low flow level. Decided to exchange building valves of Cyclo manifolds because we can't close one of valve and inlet diameter of this valves smaller than pipes that can reduce water flow. | Andriy Lytovchenko | 01/16/20 5:12 PM
Row 70 | with Heyuan. Connectors of SM water cooling exchanged. Water opened at Room1, need check Flow switches adjustment of magnets Q1N1,Q2N1. Also water runs through GTR2 and GTR3 without Q1N2,Q2N2,Q1N3,Q2N3 and Scanning magnets. Going to exchange connectors of Scanning Magnets at GTR2,3. Even water runs trough all rooms Flow switch of General Cooling show Zero flow. After finish all connections at rooms we need investigate problem with main flow switch, clean strain of General cooling. | Andriy Lytovchenko | 01/13/20 5:12 PM
Row 70 | With Zihua. Found leak at GTR1 from scanning magnets hose connection, also 2 other magnets have problem with tubing. GTR2 hasn't issues, GTR3 impossible to check as it on 180 Degree. Refill Tank | Andriy Lytovchenko | 11/25/19 3:42 PM
Row 70 | We found general cooling OFF & empty. We refilled the tank but the refill water flow rate is very very low. I should take a lot of time to refill general cooling. | Pascal PELERIN | 11/23/19 12:32 PM
Row 70 | We worked on the general cooling pressure regulation. We found that the signal from the pressure probe is not stable. There is a 300Hz signal sometimes around the DC (average) value. We set a 100µF/50V capacitor across the pin X & B (pressure gauge feedback signal & ground). If the configuration of the parameters are : pressure regulation - proportional gain = 5%/bar & Pressure regulation - integrative time = 1s, it works very fine. In this case, the system oscillates a little bit during & after the ramp)up, but, the oscillations decreases & the regulator reaches a steady status (frequency is stable by 0.1Hz). The screen shot have been added in the logbook. | Pascal PELERIN | 11/18/19 12:14 PM
Row 70 | We were not able to start up the general cooling because the parameter "pressure regulator - proportional gain" set to 1% instead of 10%. It caused the general cooling start-up to be very slow. The general cooling flow eletta did not switch ON before the end of the Archestra delay. We change this parameter to 5%/bar. Pressure is not stable, motor pump is not constant. We should check & clean the pressure gauge. Maybe the pressure gauge signal is noisy by the general cooling frequency regulator. Now, pump 1 can be switch ON under frequency control & pressure control. | Pascal PELERIN | 11/14/19 10:42 PM
Row 70 | 5 Nov 2019, Tuesday, General Cooling blocking today, mid-morning. Failed Eletta flow meter response identified and repaired. Site may need to consider ordering a replacement to have on-hand. | Mike Welch | 11/06/19 3:27 AM
Row 70 | New probe (63.07.02.0005) has been ordered by Chenchen | Pascal PELERIN | 10/30/19 9:43 AM
Row 70 | issue with the MC conditioner resistivity probe: we suspect this probe to be damaged because resistivity is still high with a new deionisation bootle. | Pascal PELERIN | 10/30/19 9:34 AM
Row 70 | issue with low pressure, need investigate driver or Arkestra settings | Andriy Lytovchenko | 10/17/19 6:03 PM
Row 70 | - Beam line magnets flushed; - General cooling connected with deionization bottle, flow rate and pressure adjusted; | Andriy Lytovchenko | 10/17/19 6:02 PM
Row 70 | MPSS cabinets connected to water manifold, flushed no leakage. beamline magnets flashing ongoing | Andriy Lytovchenko | 10/16/19 12:45 PM
Row 70 | MPSS water manifold assembled, connection to MPSS cabinets ongoing. Cyclo water manifold connected, beam line magnets manifold connected. going to flash all lines and after connect deionization bottles to General cooling group and test. | Andriy Lytovchenko | 10/15/19 12:19 PM
Row 70 | We have to flush again the water, still pump running before cleaning. | Rene BREEMANS | 10/02/19 1:53 PM
Row 70 | Back on site and looking for the problem: main breaker of the pumps power supply was not setted to the right protection value. Set to 90Amps and pumps are working. | Rene BREEMANS | 10/02/19 1:50 PM
Row 70 | Try to make the pumps working with Achestra: nok. No alarms but nothing happen after turn on the pump. Pumps not running. Checking with Andrew Ray why it's not working: issue with the altivar settings as the control cabinet is full with 2 altivar in profibus. Altivar set with new config: still not ok with archestra and pumps are turning in the wrong direction and epf1 alarm on the altivar of pump1. Solving issue for both but Andrew had to work on the config of the altivar. Solved but pump are off after a few minutes. Decide to let Andrew looking what's wrong during our night. | Rene BREEMANS | 10/02/19 1:48 PM
Row 70 | We managed to make the customer understand how to flush their building cooling. Their water was black, it will take 1 week to clear. | Jean-Michel Denis | 09/21/19 10:02 AM
Row 70 | We set the Altivar parameter. We started both pump. Pump 1 and 2 are inverted. They also were both turning. | Jean-Michel Denis | 09/21/19 9:53 AM
Row 70 | Did the bobla cabling this afternoon | Jean-Michel Denis | 09/19/19 4:49 PM
Row 71 | 15 Nov 2019, morning shift, with Chenchen and GZ Team. The Deflector flow indicated is NOK, low flow at 1.6x0.6. A setting of about 1.5x0.6 =OK. This needs further research and flow verification measurement, at this setting, for review. General Cooling is set at 8.5 bar (pressure control, but flow is changing/cycling up/dwn some. | Mike Welch | 11/15/19 9:52 AM
Row 71 | We checked the threshold of each eletta of the cyclo manifold. Settings are correct. | Pascal PELERIN | 11/14/19 10:37 PM
Row 71 | work with CCO, finish the cyclo manifold UT | Shuge Miao | 08/22/19 8:39 AM
Row 72 | Chiller run, temperature of ODP cooling adjusted, procedure passed. | Andriy Lytovchenko | 09/26/19 1:43 PM
Row 72 | Problem on customer chiller side. pressure is not enough for cool down our system. flushing and cleaning chiller side. | Andriy Lytovchenko | 09/19/19 5:04 PM
Row 72 | Connected water manifold, flow meter connected, after stop leak tight manifold pipes, run procedure, passed at 90%, As a customer not ready run Chiller, we couldn't adjust temperature yet. | Andriy Lytovchenko | 09/06/19 4:37 PM
Row 73 | 14 Nov 2019, morning shift with Chenchen and Andriy. Following prior afternoon shift work, (13 Nov) soldering of upper H3 DEE performed by Pascal and Shuge, base vacuum reached, 3.1e-6 by this morning (penning2). We performed a leak check, as in the past, by applying Air Pressure (6 bar) on H3 cooling line, no change in vacuum noted. Leak is finally repaired! | Mike Welch | 11/14/19 12:23 PM
Row 73 | 5-9 Nov 2019: Continued DEE/Pillar O-ring swaps, including re-surfacing O-ring contact faces on Pillars and DEE, still see small leak in the e-6 range when Air pressure is applied to H3 cooling piping. Setup to isolate the DEE Flow tube from the O-ring and Pillar interface for Vacuum leak checking for O-ring leak check, separating DEE from cooling circuit while installed normally. Performed leak check in this configuration and found there is no leak in the O-ring DEE/Pillar/cooling tube interface surfaces. Next shift, need to check DEE for leaks. After isolating Dee from O-ring, Pillars, cooling tubes, now we tested the DEE, by removing DEE cooling blocking. Noted deformation of hole plug, indicating increased vacuum on the DEE side. All tests results lead to suspect that leak is in DEE. Attempts were made to pinpoint the leak with DEE under vacuum, in the past weeks and again this week, but leak seems to seal under vacuum and only appears when under pressure. Therefore, the DEE needs to be pressurized to find the leak, with the help of a Sniffer tool attachment for the Leak Detector. The DEE in question is currently outside of the Cyclo. | Mike Welch | 11/12/19 5:52 PM
| Updated times 32 lost, existing, 32 added to total time, updated. | Mike Welch | 11/14/19 6:46 AM
Row 73 | 11 Nov 2019-Setup DEE for testing with pressurized Helium check and Sniffer detection. Found O-ring mating surface new scratches from prior setup. Next time, avoid use of any metal in the DEE O-ring test setup. re-surfaced O-ring matting surfaces. Sniffer testing: Confirmed leak is at small pillar side, DEE Tubing brazing. Note: Found site owned Sniffer. Someone onsite recognized it. Need a Torch and gas to re-flow or add solder in this area. None available, on-site. Requested resource or equipment to repair solder joint of DEE cooling tube end. | Mike Welch | 11/12/19 6:26 PM
Row 73 | We add the beam line pipes between PMQ & first extraction quadruple. We also set the caps at both pillars H3 upper. | Pascal PELERIN | 10/31/19 4:49 PM
Row 73 | We leak check the dee from inside cyclo: we have an helium response on both side of small pillar. We removed the dee for testing the dee only. We found no leak on the dee. Thus, the leak should be due to the groove in the pillar ( deeper than expected). The O-ring is not compressed as expected. Chenchen already order bigger O-rings. | Pascal PELERIN | 10/30/19 5:09 PM
Row 73 | We leak check the dee from inside cyclo: we have an helium response on both side of small pillar. We removed the dee for testing the dee only. We found no leak on the dee. Thus, the leak should be due to the groove in the pillar ( deeper than expected). The O-ring is not compressed as expected. Chenchen already order bigger O-rings. | Pascal PELERIN | 10/30/19 5:09 PM
Row 73 | Work with Shuge, get the adapter from chenchen, try to connect it to leak detector, but still to much water in circuit, connect the small source pump on it to pump out the water whole night. | Lipeng Han | 10/29/19 5:02 PM
Row 73 | Work with Shuge, remove the H3 upper Dee and pillars, O-ring look ok, put back the pillars, take two pixel glass to seal the pillars water circuit then restart vacuum to see if we still have leak | Lipeng Han | 10/25/19 1:14 PM
Row 73 | work woth Shuge, due to the leak at Upper dee&pillars H3, we test circuit with helium, same result, decide to open cyclo to check cavity screws, foudn upper H3, upper/lower H1all have screws not tight enough, after tighten them, close cyclo, restart vacuum, will check vacuum level tomorrow. Also chenchen help to order the fitting to cavity cooling circuit to do the leak check tomorrow. | Lipeng Han | 10/23/19 4:11 PM
Row 73 | We installed a bunch of pipes at the extraction to be able to perform future leak test more easily | Jean-Michel Denis | 09/17/19 5:06 PM
Row 73 | Finish the Piranis connector. Piranis calibrated. vacuum in source chamber and in cyclo can be done via archestra MID not passing right now: 6.2.8 steps not passing, we will try again on monday by pumping more time. | Jean-Michel Denis | 09/06/19 4:53 PM
Row 73 | Time lost on looking for Piranis connector, found part of them in the storage room the other part in the vault ... Time lost on making Penning cables | Jean-Michel Denis | 09/05/19 5:10 PM
Row 73 | 04/09/19 @jean-michel.denis@iba-group.com , @mohan.naikare@iba-group.com , @shuge.miao@iba-group.com Started with a leak detection. The vacuum pressure doesn’t go below 7e-1mbar with the roughing pump, so we expect a huge leak in the order of 5e-6mbar/5e-5mbar We found a leak on the Source VAT valve. From the outside of the Valve, not from the source chamber. The leak was measured at 8e-5mbar.l/s | Jean-Michel Denis | 09/04/19 11:47 AM
Row 73 | turned on primary pump, but seems there's some leak, can only reach 8.x e-1 mbar | Shuge Miao | 09/02/19 4:55 PM
Row 74 | MID done and passed: air pressure 6,1 mbar at cabinet inside the vault. | Rene BREEMANS | 10/01/19 9:39 AM
Row 75 | Procedure fulfilled & passed by PP on Nov11th | Pascal PELERIN | 11/29/19 2:01 PM
Row 75 | Note: no hours are entered for line 75, need to determine correct value, in hrs used. | Mike Welch | 11/15/19 9:57 AM
Row 75 | work with CCO, finish the main coil manifold UT | Shuge Miao | 08/22/19 8:39 AM
Row 76 | mobile capacitors left and right side are broken, no spare parts, ask Janet to order two of it. | Lipeng Han | 10/23/19 4:12 PM
Row 76 | MID finished | Jean-Michel Denis | 09/11/19 4:38 PM
Row 76 | everything measured, MID nearly done. Still the measurement inside the cyclo to perform We lost a bit of time because the tachometer of the mobile capacitor motor was connected backward. | Jean-Michel Denis | 09/10/19 5:38 PM
Row 77 | redo the front page. | Sherman Chang | 01/17/20 9:11 AM
Row 77 | MID finished | Jean-Michel Denis | 09/11/19 4:38 PM
Row 77 | Nearly everything was measured. We had to readjust the 3 cavity pick up as they were out of tolerances. still the Spinner line pick up to measure/calibrate. | Jean-Michel Denis | 09/10/19 5:39 PM
Row 77 | We had to adjusted the coupler-Type N adaptor to fit at the coupler side. Took some RF measurement without the source, all cavity pick ups seems out of tolerance. | Jean-Michel Denis | 09/04/19 5:07 PM
Row 78 | Procedure fulfilled & signed. On the Mikhail desk. | Pascal PELERIN | 11/13/19 5:18 PM
Row 78 | FPA on dummy load test done. Test passed. See \\T30gzptcfs1\Public\Cyclo\installation\Calibration\RF\FPA\FPA ON load_at_Site. We also fix the IMPA FWD & RFL power monitoring problem under Archestra ( wiring twisted in IMPA). We calibrated the IMPA FWD & RFL power. We also calibrated the SSA FWD & RFL power. Start to the fulfill of procedure. | Pascal PELERIN | 11/12/19 10:14 PM
Row 78 | test FPA until 100kW. record the parameters / co;pare to lln. we found better efficiency, better gain, less screen current. suspect calibration of detector box being wrong.we checked it and found wrong calibration (95kw instead of 88 - 90kw). we correct the calibration, restest fpa. parameters are now closer to fpa parameters. we saw that impa fwd and rfl power are twisted under archestra. | Shuge Miao | 11/11/19 11:06 PM
Row 78 | test FPA until 100kW. record the parameters / co;pare to lln. we found better efficiency, better gain, less screen current. suspect calibration of detector box being wrong.we checked it and found wrong calibration (95kw instead of 88 - 90kw). we correct the calibration, restest fpa. parameters are now closer to fpa parameters. we saw that impa fwd and rfl power are twisted under archestra. | Shuge Miao | 11/11/19 11:06 PM
Row 78 | On shift with Shuge. Tetrode conditioning delayed due to General Cooling fault. Tetrode conditioning started today, 1651 Watts at full power. | Mike Welch | 11/06/19 12:12 AM
Row 78 | On shift with Shuge. The new Tetrode, replacing the existing (2015) Tetrode was tested under HV for shorts or other faults. All stand alone tests Passed. The new Tetrode was installed and is ready for conditioning. Thanks to Pierre C. (joined site today) for assisting in finding Tetrode cooling flow problem. | Mike Welch | 11/04/19 8:17 PM
Row 78 | On shift with Pascal and Shuge. Entered Team time for 31 Oct 19, 2 hrs, IMPA I/O power section. FPA section pending receipt of replacement Tetrode. 2 hrs entered for dummy load prep/setup for testing. 1 hr added for FPA roof modification. (see Task list for Vacuum assy/repair work, today). | Mike Welch | 10/31/19 11:40 PM
Row 78 | Dummy load has been connected to the FPA output. Heliax has been connected to FPA input. The roof has been modified for the Heliax. | Pascal PELERIN | 10/31/19 4:48 PM
Row 78 | IMPA has been tuned on Zhuozhou dummy load. See \\T30gzptcfs1\Public\Cyclo\installation\Calibration\RF\IMPA. | Pascal PELERIN | 10/31/19 4:47 PM
Row 78 | tetrode of FPA offically die, wait fro the new one.dummy load over heat due to lost of cooling, need to borrow from Zhuozhou | Lipeng Han | 10/23/19 4:55 PM
Row 78 | FPA tetrode tested by HVPS. Test failed. THe tetrode does not hold the expected voltage. We think it is due to a tiny leak, causing vacuum to be worst than expected. We will run filament at full power for 1 or 2 days, expecting the recovering of the vacuum after that. Maybe tetrode will fail during the FPA tuning. | Pascal PELERIN | 10/18/19 5:05 PM
Row 78 | IMPA tetrode test done. Test passed. Tetrode has been installed in IMPA. IMPA has been connected to the 125kW dummy load. The dummy load has been piped to general cooling. | Pascal PELERIN | 10/18/19 5:02 PM
Row 78 | SSA test done. Test passed. | Pascal PELERIN | 10/18/19 5:01 PM
Row 78 | Power pick ups tuned. Spinner Line Resonance measured. | Jean-Michel Denis | 09/12/19 5:17 PM
Row 79 | Reviewed by CCB T30: xwchang, ccyuan. OK to close and set to QA for validation. | Sherman Chang | 07/23/20 10:22 AM
Row 79 | May 25 RGA data recording; | Shuge Miao | 06/03/20 9:58 AM
Row 79 | we add the tube / extension for setting RGA without HC connector collision. We should be able to make a RGA analysis now | Pascal PELERIN | 10/30/19 5:06 PM
Row 79 | try connect RGA on, but seems there's no room for it, the HC connector block it's way. need to add the extension tube next time | Shuge Miao | 10/17/19 5:13 PM
Row 79 | Waiting for MC coil to launch RGA test. | Rene BREEMANS | 10/01/19 9:40 AM
Row 79 | RGA sensor and flange installed on site. High Vacuum launch for timing measurement. | Rene BREEMANS | 09/27/19 11:13 AM
Row 79 | Have to do the RGA measurement | Rene BREEMANS | 09/27/19 9:25 AM
Row 79 | After deeper investigation on the customer cooling flow (found a flow meter) 128M³/h on a room we decide to go ahead with the cooling groups. ODP vacuum cooling group is working at nominal value. We launch the primary pump and we achieve good base vacuum on the ODP. We were in the impossibilty to launch the high vacuum due to bad feedbak of the odp piranni to the ACU. After investigation with Andrew Ray, the signal on I5,7 of ACU was not send from TPG300 to ACU as the DB15 cables were missing. 3 cables were installed (TPG Cyclo, source and from the CC PS) successfully. After that we were able to launch the high vacuum in the cyclo: reaching 2,6E-6mbar) | Rene BREEMANS | 09/27/19 9:24 AM
Row 80 | 2020-7-15 Done with test, still have paperwork to do | Shuling Peng | 07/15/20 6:06 PM
Row 80 | We have to receive Validated test report from subcontractor before this test | Andriy Lytovchenko | 12/31/19 11:19 AM
Row 81 | @shuling.peng@iba-group.com Please check the bug comments and update status. Thanks! | Sherman Chang | 07/21/20 3:24 PM
Row 81 | DEC 7th, Done with UPS test | Shuling Peng | 01/07/20 11:31 AM
Row 82 | Reviewed by CCB T30: xwchang, ccyuan. OK to close and set to QA for validation. | Sherman Chang | 07/22/20 10:15 AM
Row 82 | MID nearly done. The polarization of the extraction quad and steering need to be checked when we will have beam on degrader. | Jean-Michel Denis | 01/23/20 12:07 PM
Row 82 | Procedure 7493 - MC part - has been done. No problem found, everything works as expected. | Pascal PELERIN | 11/19/19 9:06 AM
Row 82 | MCPS modified: we add the 100µF & 1500µF capacitors on the entrance transformer current measurement system (to avoid Zhuozhou downtime). | Pascal PELERIN | 10/28/19 5:32 PM
Row 82 | Yoke temperature electronic unit done. Need to check that calibration is ok. | Pascal PELERIN | 10/28/19 5:30 PM
Row 82 | After replacing AL2 auxiliary PS by adjustable PS from tools rooms, the problem disappears. We order new ALE1502D pour replacing AL2 inside MCPS | Pascal PELERIN | 10/24/19 9:40 AM
Row 82 | Play procedure 7493. MCPS can not reach 720A. Up to 700A, MCPS stop by itself with "ground" alarm. We investigated the problem & found PS output oscillating up to 700A. DC voltage is stable - not depending on MC current. AC voltage between phase is stable - not depending on MC current. AC voltage between ground & phase is stable - not depending on MC current. We found only MC output voltage oscillating up to 700A. | Pascal PELERIN | 10/10/19 5:14 PM
Row 82 | CCPS ok. Found polarity wrong. Correct it. CCPS calibration done. See logbook. | Pascal PELERIN | 10/09/19 5:07 PM
Row 82 | MCPS ok in local & remote mode. Found power dissipation 3x bigger @600A at the junction pancake 4 <=> Pancake 5. | Pascal PELERIN | 10/09/19 5:06 PM
Row 82 | add 1 connector (61.03.03.0153) at the HV line, deflector side. | Pascal PELERIN | 10/09/19 5:04 PM
Row 82 | Miao finished the Harmonic coil section of the MID. Feedback in archestra recalibrated | Jean-Michel Denis | 09/17/19 8:54 AM
Row 82 | Tried to start the Harmonic coils, it was possible in local but not in remote. We were having a feedback of internal fault in Archestra but no fault on the PS. We troubleshooted the Harmonics PS. Seems like the display of the HC1-3 is not stable. 4 wires in the cables from ACU to HCPS were inverted (WB82-B0A93-B0A95) pin K&L and pin D&E After correcting it, the remote mode is working properly | Jean-Michel Denis | 09/11/19 4:41 PM
Row 83 | after complete the procedure, restored RF condition. | Shuge Miao | 11/22/19 10:44 AM
Row 83 | complete RF polarization test (7494) successfully. | Shuge Miao | 11/22/19 10:44 AM
Row 83 | We continued the RF conditioning. We tuned phase loop (we found cavity tuner tacho polarity being wrong). We got 55kW continuous. System is ready for playing the procedure. | Pascal PELERIN | 11/21/19 9:35 PM
Row 83 | 21-Nov-2019, Thurs. On shift with Shuge-Continued RF Conditioning process, Pulse mode waveshape approaching good condition. Manually reached 18 KW fwd power output, rfl power output 0.05kw, 40kv setting in LLRF. LLRF settings and auto mode needs additional work. end of shift. We did increase the Yoke Heater set-point from 35 to 37C, to see if Yoke temp may have an impact on the RF pulse result. By the end of shift, the pulse shape was acceptable. It is not know if Yoke temp difference/change had an effect or not. | Mike Welch | 11/21/19 2:52 PM
Row 83 | We start the RF conditioning with a short-circuit at LLRF TSS, radial probe FULL IN & radial probe wires disconnected. | Pascal PELERIN | 11/20/19 7:33 PM
Row 84 | Procedure has been done. No issue. Procedure signed & left on Mikhail desk. | Pascal PELERIN | 11/28/19 5:53 PM
Row 84 | 19 Nov 2019, Tues, adding Information from afternoon shift report. Source #3: 2.4mm set point  2.4mm in the cyclo. Chimney is // to the puller We tuned the full IN LS (1.8mm) & FULL out LS (7.1mm). Service position is now 6.5mm We extracted & insert the source / checking that there is no mechanical collision between the chimney & the puller: ok no collision, there is a ~2mm gap between the chimney & the puller. Slits level is a little bit too low, but, good enough for first beam test. Calibration parameters: ACU – Analog IN 1 / Source azimuthal movement positioning monitoring: Slope = 0.00023 Offset = 3.7 ACU – Analog OUT / Source azimuthal movement position Slope = 4226.64 Offset = -15570 | Mike Welch | 11/21/19 7:50 PM
Row 84 | 19 Nov 2019, Tues, on shift with Shuge. Due to questions about valid Archestra parameters, started Source Puller gap check. Found gap was different than expected, due to invalid parameter set loaded. Did not finish checking Source Puller gap check. There was some trouble moving the Source IN or Out and this problem spanned shift change. It was corrected by the end of the day, afternoon shift. Source Calibration is done for Source #3. There are only 2 Chimneys on-site, Source#2 also has a Chimney. Source #1 has no Chimney. Note: There are a total of 3 Filaments available, 1 is installed in Source#3 which also has a Chimney. It is Ready for use and calibration has been re-done, by afternoon shift, due to lost parameters. Historical data indicates that all 3 delivered Sources measure equally for Source Puller gap. | Mike Welch | 11/21/19 7:25 PM
Row 84 | 18 Nov 2019, Monday, on morning shift with Shuge. • The Base vacuum is 1.2/1.3 e-6 • Leak checked all the 3 sources. • Checked each Source Arc insulation with Megger. All 3 sources are 50Gohms or higher. • Fixed the leaks found. The Hydrogen line to the Sources on the connection stack was the only leak found and was repaired by tightening a small amount on the 1/16 line connection. • Source #3 has a Chimney and Filament. There are only 2 more Filaments available. The site needs more replacement chimney & filaments. • Mid 7495 Source Injection IT, some steps performed leading up to Arc PS test. o Step 7.4.4, prescribed 4 turns on Hydrogen Porter valve only required 2 turns, at 3.4 bar for 2e-5 vacuum/pressure. o Checked the continuity and polarity of the ARC PS Twinax cable to the ISPS in PSR. Each end of this cable (contains WB62) is wired the same and appears to be correct. We questioned this, because procedure MID 7495 ARC PS results show a negative current and Voltage when programmed. May require additional setup at BCREU, etc. (to be continued, later. Not critical for the current week mission). • With the Source #3 inserted and the Hydrogen turns at zero, the base pressure is now 1.2/1.4 e-6. Compared to this morning, the vacuum/pressure is about the same, with the Source inserted, +0.1 e-6. Source leaks resolved. Note: 2 New Hydrogen bottles are delivered and 1 is installed anew, as of end of last week.. | Mike Welch | 11/18/19 5:29 PM
Row 84 | 15 Nov 2019, morning shift. Reviewing procedure, setting the H2 pressure threshold to the lower value of, 1.5 bar (on decreasing pressure) for interlock, then set to the operating pressure for, 3.4 bar. This is finished. Also, a small leak was found in the control air pressure line at the Air solenoid in the H2 box, and is repaired now. | Mike Welch | 11/15/19 7:39 AM
Row 84 | We found the H2 pressure switch OFF although the H2 line is pressurized. We change the threshold to get the signal OK. Need to check that the threshold is correct versus the procedure. We tried to start up the filament PS remotely. We found pin C<=>D & K<=>L in verted at the 23 pins burndy connector of filament PS. We did not correct the wiring because we did not find the burndy pin extractor tools. After correcting the synchro signal & Reg out signal at the BCREU rear side, we got the arc PS ON. We calibrated the arc current, arc voltage is more or less calibrated. To be fine tuned when source is producing beam. | Pascal PELERIN | 11/14/19 10:45 PM
Row 84 | 14 Nov 2019, On shift with Chenchen and Andriy. (1.) Source#3 was prepared with a Filament and Chimney and inserted, following final Vacuum system verification steps (H3 DEE cooling/vac leak repaired now). (2.) There were some cooling system problems/variation in stability. MCPS ready in set to remote. (3.) Removed the H2 OK bypass and continued. H2 pressure sensing checked and is near 3.4 bar. Reviewed for most recent version setting in bar. All Source Injection interlocks cleared without bypasses, except for a TSS interlock. Need to explore how to clear this interlock or why it cannot be cleared, yet. | Mike Welch | 11/14/19 1:02 PM
Row 84 | Some cabling was finished last week. Verified all cable routing is done correctly. System is electrically ready. Set-up a dummy load provided by Pascal for the ARC PS test. | Mike Welch | 11/12/19 7:03 PM
Row 84 | insert source,switch on filament, no issues. Now wait for hydrogen | Lipeng Han | 10/29/19 7:55 AM
Row 84 | Someone misconnected the cable WB86(from beckhoff to ACU) to a spare one port, correct it. | Lipeng Han | 10/28/19 12:41 PM
Row 84 | We adjusted the source position and calibrated the azimuthal movement. Puller parallelism, angle and Horizontal and vertical slit-puller position adjusted Calib azimuthal done on source 1 and Checked on S2 and S3 Archestra SP to have 2.4mm in cyclo: S1: 2.4mm S2: 2.45mm S3: 2.4mm | Jean-Michel Denis | 09/10/19 9:30 AM
Row 84 | Powered on the beckhoff drivers and adjusting the source movement switches Powered on the DC driver rack, Azimutal movement is working properly. | Jean-Michel Denis | 09/05/19 5:13 PM
Row 85 | Procedure completed, signed & set on Mikhail desk | Pascal PELERIN | 11/27/19 3:21 PM
Row 85 | Work with Pascal, put tetrode inside FPA cabinet, set filament on to condition this tetrode | Lipeng Han | 10/22/19 9:51 PM
Row 86 | July 6 finish MID-7525 part 7.7.2, 7.8, 7.9, 7.10 | Shuge Miao | 07/10/20 9:24 AM
Row 86 | July 3 7.9.4 & 7.9.5 passed; at lower beam current, extraction efficiency still too low. | Shuge Miao | 07/03/20 4:02 PM
Row 86 | 2020-Jun-03 1.Switch off main coil power supply due to customer side strainer dirty. | Haiyuan Pan | 06/03/20 1:27 PM
Row 86 | June 2nd record shims inside cyclo | Shuge Miao | 06/03/20 9:55 AM
Row 86 | Chenchen remove the S1E and wait degrader to cool down, we can not shoot beam, so we replace the potential meter of YLCU, did MID7525 no beam part. Shuge heat up the RGA for the night. | Lipeng Han | 05/20/20 8:11 AM
Row 86 | After restart RF and deflector, start to do CC sweep, MC sweep, deflector voltage& exit position sweep, deflector shadowing, HC sweep | Lipeng Han | 05/03/20 5:48 PM
Row 86 | customer cooling system down, which cause RF down and vacuum down... | Chenchen Yuan | 03/29/20 4:13 PM
| 2020 03 25 morning shift | Chenchen Yuan | 03/29/20 4:14 PM
Row 86 | customer air supply down, we lost vacuum, and troubleshooting with customer, air supply system need frequently maintenance ( air compressor need maintenance, filter need replace, air dryer need maintenance) , we fixed it, and recover vacuum and RF conditioning, we lost 4 hours almost. | Chenchen Yuan | 03/20/20 4:57 PM
Row 86 | We open the cyclo & reroute the collimators cables. | Pascal PELERIN | 12/02/19 12:06 PM
Row 86 | We shooted first beam on BS on Nov 28th, 2019 | Pascal PELERIN | 11/29/19 2:02 PM
Row 86 | Beam line pipe between 2d PMQ & first extraction quadrupole has been removed. flange has been installed instead. BS has been set in the cyclo. | Pascal PELERIN | 11/19/19 9:04 AM
Row 87 | Reviewed by CCB T30: xwchang, ccyuan. OK to close and set to QA for validation. | Sherman Chang | 07/22/20 10:16 AM
Row 87 | Big issue with the profibus of the MCPS. Adresses are good, 5 for MCPS and 22 for PAC 3200. No communication wtih the ACU is working. Impossible to access the back of the iocaste. | Rene BREEMANS | 10/04/19 1:01 PM
Row 87 | Main breaker of the mcps powered on: all ok. | Rene BREEMANS | 10/02/19 2:44 PM
Row 87 | Checking the settings to put the MCps in service: ok for main transfo and the switch selector. Customer hook-up powered on and tension checked 412v and 238v: ok selection line hook-up actuated and automatic line switch select the powered circuit. Hook-up client powered on and ok. | Rene BREEMANS | 10/02/19 2:43 PM
Row 87 | Checked the cooling circuit: ok. | Rene BREEMANS | 10/02/19 2:38 PM
Row 87 | Leak fixed and circuit under pressure during few minutes before leaving the site: no more leak. | Rene BREEMANS | 10/01/19 9:14 AM
Row 87 | both pumps are cabled and ready to work. 2 phases of the pump 2 were inverted: corrected. A tiny leak appear inside the MC PS at the souple pipe hose. To be fixed. | Rene BREEMANS | 09/30/19 1:11 PM
Row 87 | New relief valve installed on the mcps piping by customer. | Rene BREEMANS | 09/30/19 10:50 AM
Row 87 | New relief valve has to be installed as the installed one is opening at 25 bars. Parts ordered on thursday 26/09 and arrived the 30/09. Relief valve tested at 8,1 bar before installation. | Rene BREEMANS | 09/30/19 10:50 AM
Row 87 | All flexible pipes are connected for he cooling of the MC PS: ps side and cooling group side. | Rene BREEMANS | 09/27/19 2:18 PM
Row 87 | Flexible pipes for the MC ps has been installed by Better in the PSR between the buildings and the internal valve. Missing a good relief valve with 8 bars release pressure. To be installed on monday 30/09. Connection between the group and the building have to be redo (wrong holding cap). | Rene BREEMANS | 09/27/19 11:49 AM
Row 87 | Tested group with 2 as pump have no cable. Pressure is 17,2 bar and flow 110L/min on eletta. | Rene BREEMANS | 09/27/19 11:22 AM
Row 87 | Yesterday, we worked on the chiller issue. After troubleshooting, the flow rate issue was due to insufficient pressure at the output of the chiller pumps | Jean-Michel Denis | 09/19/19 9:19 AM
Row 88 | Reviewed by CCB T30: xwchang, ccyuan. OK to close and set to QA for validation. | Sherman Chang | 07/22/20 10:16 AM
Row 88 | @shuge.miao@iba-group.com , please add the comments and print email, attached into the site test BUG, should be check by QA. Thanks! | Sherman Chang | 07/16/20 11:50 AM
| email printed and attached to test report. | Shuge Miao | 07/16/20 3:51 PM
Row 88 | June 1st restore parameters in YLS driver, as the flying engineer says we should not modify the parameters although the MID asks to do so. continue YLS test section 7.5 | Shuge Miao | 06/03/20 9:53 AM
Row 88 | May 29 play MID 37919; correct wiring issue inside cabinet; potentiometer calibrated. | Shuge Miao | 06/03/20 9:51 AM
Row 88 | May 27 troubleshooting YLS issue, correct the wiring; | Shuge Miao | 06/03/20 9:49 AM
Row 88 | Open cyclo to install the vacuum pipe between PMQ and Q1C, then try to do the YLS UT, but we found out the 10 V PS inside cabinet doesn`t work properly, the reading of those two jacks position is jumping around, order a new one locally, try to use a external PS, but still the reading is jump too much to calibrate it. After we close the cyclo, vacuum stuck at 2.6mbar. We believe the vacuum pipe was not installing properly, will open cyclo tomorrow to check. | Lipeng Han | 05/12/20 5:31 PM
Row 88 | Mar 4th 1. check YLCU status, found no CF card inside. 2. found 2 CF card in office, but none can work. The original one has been overwritten. | Shuge Miao | 03/04/20 4:34 PM
Row 88 | We have some issue to connect to the YLCU. At first we were able to ping it, but not connect to it via VNC. So we burned the card again. But now we cannot ping it anymore. Waiting for Andrew support to help the troubleshooting. | Jean-Michel Denis | 01/23/20 11:59 AM
Row 89 | Page need to be update, please handover to Chenchen after change done. | Sherman Chang | 07/22/20 10:17 AM
Row 89 | When we are at §7.1.17 to modify the default RF frequency, we could not start RF as TPO said, but we could not see red error on Archestra, it`s in grey. Ask Andrew to modify it. All finish except above with new LLRF 4.27 NI4003_1950_10139 | Lipeng Han | 05/20/20 2:24 PM
Row 89 | Procedure nearly done for LLRF NI4003-3818-003 from zhouzhou. But failed the RF detection glitch. So we lost a shift trying to do the MID as this LLRF is defective and need to be sent back to LLN | Jean-Michel Denis | 01/25/20 3:14 PM
Row 89 | Procedure done only for Mod power limit / Dee voltage gain / Vreg loop gain adjustment because LLRF is V4.1, we need V4.27 for CFDA. | Pascal PELERIN | 11/29/19 2:04 PM
Row 90 | equipment arrived from LLN, need a big time slot to do it. | Sherman Chang | 07/20/20 9:49 AM
Row 90 | need customer to prepare the equipment. | Shuge Miao | 12/27/19 3:40 PM
Row 91 | measured the temperature until the yoke got cold. complete this procedure. | Shuge Miao | 12/06/19 7:41 PM
Row 91 | 20 Nov 2019, Wed, on shift with Shuge. Shuge performed site BUG 81876 for Earth Leakage Current Breaker update, 300mA to 30mA. Added 1 hr in hours used for this. Shuge to updated Bug. | Mike Welch | 11/21/19 7:37 PM
Row 91 | 15 Nov 2019, morning shift with Chenchen. The IVVP-006 6, the Cyclotron temperature is stable, 35 C set, in range of, 34.72-34.85, over 2 hr period. | Mike Welch | 11/15/19 1:35 PM
Row 91 | 13-14 Nov 2019, on shift with Chenchen. We completed all steps, except final verification of control led temperature, to follow next shift. (see notes below). Yoke Heater system testing. MID 48694 steps notes: No spare temp probe found, yet. Per Shuge, 2 are on order. Meanwhile, we were able to re-purpose the Cyclo external temp sensor. The temperature sensors on the lower Heating Plates, 5,6, read lowest of all 8 Heating Plates and are marginal for being in spec of, +/- 3 C of Yoke Temp. These are, -4.5 to 5 C different than reference temp indication. 95% finished. Added time spent. | Mike Welch | 11/14/19 5:24 PM
Row 91 | 12 Nov 2019-On shift with Chenchen-In preparation for turning on Yoke Heaters, performed Torque verification of all heater plates. Lost power due to other work in-progress and could not complete Yoke Heater start-up testing this shift. | Mike Welch | 11/12/19 6:51 PM
Row 91 | Fan orientation corrected. We checked also that the temperature probe read the temperature of the expected heater (probe wiring correct / no twist). | Pascal PELERIN | 10/30/19 9:23 AM
Row 91 | Issue with the fan drivers, cable position inversion on the B&R output: Fan and power drivers are inverted. Corrected. | Rene BREEMANS | 10/07/19 11:58 AM
Row 91 | all thermocouples are cabled following the plans. | Rene BREEMANS | 10/07/19 11:57 AM
Row 91 | Starting doing the Yoke heater Thermocouple cables | Jean-Michel Denis | 09/19/19 4:47 PM
Row 91 | Yoke Heater assembly by local team, there are holes (M4 holes for the covers) that were not drilled yet | Jean-Michel Denis | 09/17/19 8:53 AM
Row 93 | All finish except IC calibration | Lipeng Han | 05/17/20 9:40 PM
Row 93 | Try to check all connection of BCREU, also checked the PS flags. Finally found out the IC in is set to short which is in the MID20838 §8.2.1, but we are in §8.5.34, in TPO doesn`t tell to switch IC in back to run until §8.6.5, that is the mistake in the document, need to update it. Try to do IC calibration, but the source died, we change the source to #2, also found out every time when we click prepare in BCP, the Q1CQ2C values will set to 68.75 59.75,need help fro, software to change the default value. | Lipeng Han | 05/15/20 8:55 AM
Row 93 | Found out one O-ring before Q1C was off position, after correcting it the vacuum is good. Try to shoot beam on S1E and manually adjust Q1C, Q2C, take a lot of time, still the beam shape is not good enough (Q1C=19.2, Q2C=61.2). We decide to continue BCREU VT, also upgrade LLRF to 4.27. We stuck at prepare beam. Found out the in manager mode the cyclo parameter is not updated, after update the parameter to BCP, CYCLO can do the preparation, but still could not see beam on IC cyclo. Confirm the issue is BCREU itself, will check tomorrow | Lipeng Han | 05/14/20 9:24 AM
Row 93 | Work with Shuge, check all cables from/to MCR cabinet, found one cable from BLCU to BCREU was not pulled, Shuling help to make IC cyclo signal three cable connectors, also found the ESS cabinet cable drawing has some mistake, could not connect all cables to the correct electrometers, wait for answer from R&D | Lipeng Han | 10/23/19 4:06 PM
Row 94 | after many times trying, still can't pass BCREU and BIREU at same time. we can only seprately pass BCREU or BIREU. | Shuge Miao | 07/02/20 4:36 PM
Row 94 | Revocer system from shutdown, continue with IC calibration, finish BCREU part | Lipeng Han | 06/17/20 6:14 AM
Row 94 | After cooling return, we try to do focalization, found power cable on Q1C, Q2C are reversed. After correcting it, with Q1C=68, Q2C=57, beam is well focused. Found that S1E could not show correct beam current, we think that`s the same issue as Zhuozhou. Insulation inside S1E broken. Also found the P1E could not read beam current, HV cable broken. T1CX cannot regulate, always foreseen at 60A. | Lipeng Han | 05/18/20 9:14 PM
Row 94 | Source#2 could not produce enough beam, replace to Source #3 Feels the S1E could not show the correct value, maybe need to replace it. Maybe the source issue is because the S1E Water cooling from customer side went high again. Have to stop everything and wait | Lipeng Han | 05/18/20 2:42 PM
Row 94 | Source#2 could not produce enough beam, replace to Source #3 Feels the S1E could not show the correct value, maybe need to replace it. Maybe the source issue is because the S1E Water cooling from customer side went high again. Have to stop everything and wait | Lipeng Han | 05/18/20 2:42 PM
Row 94 | Received two new IC cyclo large gap, but found one surface was hit by something that not flat anymore, we install the another one. | Lipeng Han | 10/29/19 8:00 AM
Row 97 | procedure down, need paper work | Chenchen Yuan | 07/21/20 3:47 PM
Row 97 | need create relate bug | Chenchen Yuan | 07/21/20 3:37 PM
Row 97 | According to MID 75557, install the UK of R2VCU and test in GZ site, download new code and config files to R2VCU rack. During test, we found an issue about R2VCU Deport I/O POWERLINK communication error. Action 1: According to Andiry suggestions, check the POWERLINK cable pins, it should be cross wire. Better guys made a mistake; Action 2: we use the wrong version of R2VCU install instruction MID75556-A, Nicolas suggest to use an unreleased version B, which suggest to adjust the deep switch “01”; After this two actions, we success to connect the Deport IO module, and test the hall probe feedback. test finished and passed, except paper work; | Chenchen Yuan | 04/19/20 12:14 PM
Row 97 | Add MID 75556 in ML4a. From Nicolas Gerard : R2VCU system ( redundant range verifier control unit ) has been added on Proteus Plus to be compliant to 60601-2-64. | David Ringenbach | 12/21/18 8:16 PM
Row 98 | Reviewed and OK for CCB T30. @janet.huang@iba-group.com , please help scan and upload to the corresponding bug. | Sherman Chang | 07/03/20 3:40 PM
Row 98 | Group3 test finish, for the last step "TCS check", we need use more test points for "Field vs Current", then we can get more accuracy parameters. | Chenchen Yuan | 05/17/20 12:22 PM
Row 98 | site host config for Group3 finished, BCP connection was normal, feedback was correct; need wait the MPSS working, then we can Calibrate and validated Group3 function | Chenchen Yuan | 12/25/19 6:39 AM
Row 98 | With Chechen. Group3 teslameter connected to lantronix through fiber optic cable. Lantronix IP address confirmed and connected to network. Need add Goup3 host to site host config. | Andriy Lytovchenko | 12/16/19 5:54 PM
Row 99 | turbo pump remote control mode still have issue, need control team support; | Yeehim Him | 07/21/20 11:54 PM
Row 99 | turbo pump remote control mode still have issue, need control team support; | Chenchen Yuan | 05/31/20 2:39 PM
Row 99 | because we need broke up the beam line for ESS calibration, we need setup beam line vacuum again. Beam line of vault was normal, but turbo pump was not useful, need trouble shoot "speed ok" signal,,,, | Chenchen Yuan | 05/17/20 12:24 PM
Row 99 | Beam line in vault test finished, but for GTRs need install the IC1, but before beam optics in TR, we don't want to install IC1; | Chenchen Yuan | 12/25/19 6:29 AM
Row 100 | Vault and PSR water cooling checked and passed | Yeehim Him | 05/27/20 9:34 PM
Row 101 | finish test, but still need control team support about Archestra issue for BPM function, some ID issue need fixed, please check more detail from attachment | Chenchen Yuan | 12/27/19 6:47 AM
Row 101 | working with Andiry finish Beam line elements test; next step GTRx BPM test; issue: from Archestra P1B3 can be inserted and extracted, but no status feed back; checking the status light form terminal box, no issue; BLCU phenomenon: send "insert" command from Archestra, BLCU command signal (BPM TR5 P1B5 "in" command request) & (BPM TR3 P1B3 "in" command request) & (BPM TR2 P4G2/P4I2 "in" command request) led ON; Switches signals (BPM TR3 P1B3"in" switch info) led ON & (BPM TR3 P1B3 "out" switch info) led OFF; | Chenchen Yuan | 12/25/19 7:29 AM
Row 102 | finished all test, need update test report to bugzilla | Chenchen Yuan | 12/09/19 5:43 PM
Row 102 | Together with Chechen settled IP addresses of Degrader EU (b&R, lantronix), need update calibration file which we couldn't get in bugzila, after do calibration. | Andriy Lytovchenko | 12/03/19 4:38 PM
Row 103 | BPMs Electrometer P3G1,P4G1 exchanged with Electrometer SN201708052, configured and noise level successfully recorded and procedure passed. | Andriy Lytovchenko | 01/13/20 1:40 PM
Row 103 | Tested ekectrometer from GTR1 P3G1,P4G1.. impossible to set high voltage. Resetting to default setting didn't improve. We need order new one. As we haven't in stock or temporary install other one which we removed from MCR as we installed there 2 Electrometers UK. that Electrometer till 2kV but has to work as we need only 750V | Andriy Lytovchenko | 01/10/20 5:46 PM
Row 103 | Noise from all 16 beam line Electrometers logged with EMIO diagnostic. All in green except P3G1,P4G1 electometer in GTR1, need investigate cabling of it. | Andriy Lytovchenko | 01/09/20 5:36 PM
Row 103 | Swapped gantry switches between GTR1 and GTR2 as they installed wrongly. updated software/firmware for gantries Electrometers for v5. Added 2 Electrometers from GTR 1 to DHCP list. All beam line electrometers visible from BCP. Next and last step take noise level from all electrometers. | Andriy Lytovchenko | 01/07/20 6:00 PM
Row 103 | Beam line electrometrs plugged (2 network cables didn't wired correctly), reachable from network, gantries 1,2 bpm electrometers don't work because no power cables on switches, at gantry 3 switch works but it need config. | Andriy Lytovchenko | 11/29/19 7:32 PM
Row 103 | DHCP config for all beam line electrometers updated. Need plug all cables from bilfing to server switch. | Andriy Lytovchenko | 11/28/19 11:12 PM
Row 103 | update DHCP config file for beam line and gantry electrometers. need update on server with IT | Andriy Lytovchenko | 11/27/19 3:22 PM
Row 103 | Grounded BNC rack, noise coming from C230 coaxial cables. all BNC connectors ground wired to rack ground which help reduce noise and able to do offset compensation but we also need to test noise condition with Main coil On. | Andriy Lytovchenko | 11/27/19 3:20 PM
Row 103 | ACU electrometer embcm1c firmware/software downgraded from V5 to V4.8 because of communication issue. Electrometer and BNC rack removed from ESS cabinet for investigate grounding and reduce noise. Found noise issue with one of BNC coaxial cable WB286. Going to check ESS cabinet and cable trades grounding. | Andriy Lytovchenko | 11/26/19 4:30 PM
Row 103 | Io electrometers updated and configured, IoBcreu installed in MCR, IOSMPS and IORedSMPS need install and wire inside SMPS cabinet. ACU electrometer lost communication with Archestra via Modbus..ask ControlTeam for check. also we need reduce noise level by grounding this electrometer. | Andriy Lytovchenko | 11/25/19 3:35 PM
Row 103 | Beam line electrometer's software upgraded with last version, need do for other 6 electrometrs in GTR1,2,3. BAEU/BIREU checked, correct response to TSS now. | Andriy Lytovchenko | 11/21/19 11:12 PM
Row 103 | 19 Nov 2019, Tues, on shift with Shuge. With Cyclo open, checked Collimator wiring to Electrometer. Found cables requiring a change of channels. After changes, finally, L-WB290 OK, R-WB291 OK, Up-WB292 OK, Down WB293 OK, Some wiring differences were found and corrected so that Archestra reads the information correctly on graphic displays. This was verified end to end with a simulated Beam battery source, designed for this purpose. | Mike Welch | 11/21/19 6:37 PM
Row 103 | got correct software firmware version for PTS 12, upgraded BIREU,BAEU but it still need to config them for work with BCP. Other electrometers on beam line going to be update to A.5.4.59 through network. | Andriy Lytovchenko | 11/20/19 11:01 PM
Row 103 | Updated software and firmware for all BPM and BCM electrometers. Troubleshooting with ACU BCM, need create synchro chain 24 V for test correct measurement from Archestra. | Andriy Lytovchenko | 11/18/19 11:30 PM
Row 104 | 17Jan2020. T1S2,T2S2,T3S2 popup on Arkestra fixed. We can manage magnets | Andriy Lytovchenko | 01/17/20 5:12 PM
Row 104 | from Archestra STRL2, when click the t1s2 and t2s2, no pop-up screen for this magnet | Chenchen Yuan | 12/25/19 8:03 AM
Row 104 | TDSPS config finished, and finished test except the DCCT test, because the DCCT tool calibration was over due | Chenchen Yuan | 12/25/19 8:01 AM
Row 104 | Profibus connected to all modules, set addresses. possible On from Archestra. | Andriy Lytovchenko | 11/25/19 3:30 PM
Row 105 | during the stress test, Q1N3 and Q2N3 is fail because the water cooling not enough. | Yeehim Him | 07/21/20 11:55 PM
Row 105 | almost all pass, except Q3S2 resistance which is lower than reference | Yeehim Him | 07/21/20 11:53 PM
Row 105 | Polarity check done for vault. next step Rooms 3,1,2. Leakage of nozzle magnets for all GTRs fixed. need also debug thermal switches in GTR3 | Andriy Lytovchenko | 01/16/20 5:03 PM
Row 105 | Checked magnets water cooling flowmeters feedback on Arkestra beam line, 4 flowmeters wired incorrect (in serial)..need check cabling between Bopla box and BLCU. One magnet has leakage from one of plastic tubes, going to exchange it. | Andriy Lytovchenko | 12/03/19 4:44 PM
Row 105 | PCB board of Module 4 in Quad Cabinet 1 had removed and tested. one of phases hadn't contact because of bad soldering (see attached pics). soldered, tested and assembled back in to the cabinet. All power supplies work well. Next step is check all magnets interlocks with water cooling. | Andriy Lytovchenko | 12/02/19 2:58 PM
Row 105 | Dipole power supply cabinet started without issue, all modules work. Quadrupole module 4 in cabinet 1 has problem with power (AC in red led). swapped module 4 with spare one but issue was similar. Lock out Hookups, removed module 4 and back cover of cabinet 1, checked with multimeter connections between power inputs and module 4 plug, one of phase is not upper, probably problem with soldering of plug on PCB board. PCB connection board 4 difficult to remove from cabinet for continue investigation and repear. Have to do it during next week | Andriy Lytovchenko | 11/30/19 8:05 PM
Row 105 | 40 A breakers exchanged. | Andriy Lytovchenko | 11/25/19 3:28 PM
Row 105 | 3 power supply modules exchanged. We have 2 in spare only, need order more. Also We have to get new 40A breakers for exchange them. After we will start cabinets for test. | Andriy Lytovchenko | 11/21/19 11:15 PM
Row 105 | Got 400V and we could ON Quad and Dipole PS's. 3 modules exploded. We need exchange them together with breakers and try again. | Andriy Lytovchenko | 11/20/19 11:04 PM
Row 105 | BUG for profibus upgrade resolved. We haven't 400V from HK3,4 need investigate with Better. Also HK3,4 impossible to ON from Arkestra through BLCU. | Andriy Lytovchenko | 11/20/19 12:33 AM
Row 105 | Cabinets inspected regarding wiring and dip switches setting lists. profibus configured and connected to BMS cabinet, config files of controllers updated. need setup grounding test and start cabinets. | Andriy Lytovchenko | 11/13/19 8:12 AM
Row 105 | Cabinets connected to water manifold, manifold wired and flow switch adjusted, power modules and controllers installed inside cabinets. need wiring profibus between MPSS and BLCU, adjust controllers. | Andriy Lytovchenko | 10/16/19 5:27 PM
Row 106 | sl3e frequently lost connection; Open the MSEU check sl3e wring and burndy connection, checking the wiring between MSEU and sl3e, no issue; It is very strange that after homing sl3e from BCP, "Ready" status change to green "OK", then switch to red "NO"; Once we input the distance >2mm, sl3e will move to the desired position, and status change "OK" again, and keep status green "OK" for ever; if input the distance <2mm, sl3e move to target, but status change red "NO"; | Chenchen Yuan | 12/25/19 6:54 AM
Row 106 | sl1e, and sl2e finish test; | Chenchen Yuan | 12/25/19 6:41 AM
Row 107 | q1n3 and q2n3 water flow switch threshold setting "1" | Chenchen Yuan | 06/07/20 9:58 PM
Row 107 | GTR3 magnets "water flow" & "thermal switch" interlock tested and passed; for Q1N3 there was no signal about water flow in drawings, we series connect it to Q2N3 water flow switch; it means one of water flow switch is triggered, Q1N3 & Q2N3 will shut down PS; | Chenchen Yuan | 06/04/20 11:25 PM
Row 107 | all magnets in vault are checked and passed. | Yeehim Him | 05/27/20 9:35 PM
Row 107 | straight line magnet, room 2 and 3 magnet in vault is passed | Yeehim Him | 05/26/20 11:55 PM
Row 107 | section 8, ESS magnets part finished; q8e and q10e, thermo switch and water flow switch was crossed, fixed it; test MPSS cabinets part; test TDSPS cabinets part; | Chenchen Yuan | 05/17/20 12:29 PM
Row 107 | section 7 finished | Chenchen Yuan | 12/25/19 8:04 AM
Row 107 | With Chechen. Spend time with TDSPS. Configurate settings and dip switches, wiring. Some of magnets on BTL don't allow to adjust settings from Archestra. Need check configuration of Archestra. | Andriy Lytovchenko | 12/20/19 6:09 PM
Row 107 | with Chenchen. Profibus communication test of BLCU passed. Spend time for config B&R plc for profinet name in ESS cabinet in volt. Profinet communication passed. working with Checksum part of MID | Andriy Lytovchenko | 12/19/19 3:47 PM
Row 108 | Add versions of RTCBs in override.properties, find wrong setting of SGCU, all electrometers in GTR1,2,3, Io boards became green after press select BDP. We still have to keep defined IPs of nozzles Electrometers in override.properties file for start BCP. Need investigate with IT guys | Andriy Lytovchenko | 01/09/20 5:45 PM
Row 108 | Debug BPMs wiring, have problem with wrong connection for retract, insert, movement. pneumatic connection leek, exchanged pneumatic plug . 2 BPMs don't connect to pneumatic line yet. | Andriy Lytovchenko | 12/11/19 5:11 PM
Row 108 | With Bruce, Changed settings in config files of BCP for all electrometers, SGCU, RFCU1, IoBcreu, IoSmps, IoRedSmps, BAEU, BIREU. All working except Tr2 IC2 Dosi, IoSmps because of cabling issue..need redone.etgernet cables plugs. BAEU/BIREU working on BCP, TSS bypas removed LLRF TSS signal in correct condition. | Andriy Lytovchenko | 12/06/19 2:51 PM
Row 108 | For beam line BPMs, beam stops have problem with air and wiring, last bending's flowswich for room 3 not appear on erkhestra but flowswich works itself. Need debug. cable between BMS cabinet and MCR cabinet for treatment room 0 - 8 selection pulled related to bugzila. | Andriy Lytovchenko | 12/04/19 4:06 PM
Row 108 | With Bruce, changed configurations in BCP (IPs, versions..) for SGCU, RFCU1, Io boards and nozzle electrometers.SGCU, BDCU, Io boards connected to BCP but nozzle electrometers need debugging.. | Andriy Lytovchenko | 12/04/19 3:59 PM
Row 108 | Configured Degrader EU, MCEU, IP addresses, appeared on BCP, Electrometers version need to be updated in BCP config for see them on BCP | Andriy Lytovchenko | 12/03/19 4:52 PM
Row 108 | All 12th nozzle electrometers upgraded and configured, plugged to SGCU and RFCU1. Could see feedback on LCD screens for each Gantry room. Next step see all of devices on BCP | Andriy Lytovchenko | 11/30/19 7:19 PM
Row 108 | SGCU and RFCU1 wired with Ethernet cables, gantries nozzle electrometers working but firmware/software have to be updated and configured | Andriy Lytovchenko | 11/29/19 7:35 PM
Row 108 | SGCU and RFCU1 software/firmware updated and settings configured. Ethernet cables pulled from BMS cabinet to PSR switch. Switch configured, possible to ping devices. Need to update BCP configs for software/firmware versions and see them on BCP. Next step update and plug nozzle electrometers, install io boards in SMPS cabinet. cross cable between SGCU and BLCU , bcc chain. | Andriy Lytovchenko | 11/28/19 11:25 PM
Row 109 | Field regulation for B2Gx had installed, for B1Gx, support structure had install, next step, we need install all relate accessories (hall probe, PS module...) | Chenchen Yuan | 06/07/20 10:01 PM
Row 109 | Field regulation for B2Gx had installed, for B1Gx, support structure had install, next step, we need install all relate accessories (hall probe, PS module...) | Chenchen Yuan | 06/07/20 10:00 PM
Row 109 | FIELD REGULATION POWER SUPPLY installed and wired at BMS cabinet. BUG 85637. We also need do wiring between BMS cabinet and MPSS Dipole cabinet (no clear drawing for location of connectors X1,X2 in cabinet). | Andriy Lytovchenko | 11/27/19 3:13 PM
Row 110 | working with Haiyuan, and Andiry B12E installation had finished; B1Gx installation had finished; B2Gx no enough space for installation, need support | Chenchen Yuan | 12/25/19 7:14 AM
| Field regulation for B2Gx had installed, for B1Gx, support structure had install, next step, we need install all relate accessories (hall probe, PS module...) | Chenchen Yuan | 06/07/20 10:01 PM
Row 111 | Guys, even though, we may also be better to print out WI and make "check" mark a little bit and upload to the bug to prove we done this WI. Thank you ! | Sherman Chang | 07/22/20 10:05 AM
Row 111 | MPSS installation done. no paperwork need to ccb | Yeehim Him | 07/16/20 11:03 PM
Row 111 | during installation, 3 JEMA PS modules was broken, and sent them back to LLN; during Group3 test, one JEMA PS module was broken, and sent it back to LLN. received 3 new JEMA PS modules | Chenchen Yuan | 06/07/20 10:05 PM
Row 111 | during installation, 3 JEMA PS modules was broken, and sent them back to LLN; during Group3 test, one JEMA PS module was broken, and sent it back to LLN. received 3 new JEMA PS modules | Chenchen Yuan | 06/07/20 10:05 PM
Row 111 | MPSS cabinets had installed, JEMA PS modules had installed and finished configuration | Chenchen Yuan | 06/07/20 10:03 PM
| next step, need finish test report | Chenchen Yuan | 06/07/20 10:03 PM
Row 112 | IO boards had installed by Andiry and finished configuration | Chenchen Yuan | 06/07/20 10:06 PM
Row 112 | installation had finished start SMPS IT, some issues need debgu; | Chenchen Yuan | 06/07/20 9:57 PM
Row 112 | Upgraded SMPS cabinet with two IO Boards and cables, Pyramid board removed. IO Boards upgraded with new software can communicate with BCP, SGCU/RFCU1 but show HO error. | Andriy Lytovchenko | 12/09/19 5:23 PM
Row 114 | Reviewed and OK for CCB T30. @janet.huang@iba-group.com , please help scan and upload to the corresponding bug. | Sherman Chang | 07/03/20 3:39 PM
Row 114 | daily check beam size on degrader after RF lut and ess alignment | Chenchen Yuan | 06/19/20 12:03 PM
Row 114 | we need redo and monitor the beam size on degrader, depending on yoke temperature, RF LUT...et | Chenchen Yuan | 06/18/20 1:06 PM
Row 114 | finished procedure except paper work and update bugzilla | Yeehim Him | 05/26/20 11:56 PM
Row 115 | Reviewed and OK for CCB T30. @janet.huang@iba-group.com , please help scan and upload to the corresponding bug. | Sherman Chang | 07/03/20 3:39 PM
Row 115 | test result pass; paper work finish, and send to Mikhail for reviewing | Chenchen Yuan | 07/01/20 3:54 PM
Row 115 | By Douglas 24 June 2020 We finished the rest three blocks calibration in MID3242 , the verification all passed (worst case is range error=0.023 within tolerance 0.03) . | Xinyao Li | 06/28/20 8:38 PM
Row 115 | By Mike 23 June 2020 MID 3242 is re-done with the updated Convalgo (new parameter from 3252). Before the shift end, the passthrough position, B9,8,7,6, 5 (which failed in B1234 current check last time for min range), 4 and min range of B3are passed. The BLAK file with passed range are stored in public drive. | Xinyao Li | 06/28/20 8:37 PM
Row 115 | 19 June 2020 1. Preparation: a. checked main coil tuning when the shift started. We made 2 curves and the time interval is about 20 minutes. Spreadsheet attached to this mail, but the plateau looks a little bit weird. There’s a drop of the readings on beam stop when the MC is about 753.81 A. but we didn’t use this value to generate LUT, so it won’t influence us (for now). b. then prepared system. We cannot perform ESS alignment successfully. BSS feed back showd sl3e error. Then we initialize it again (this happened several times), and it was fine. Then P1E was NOK. Reset it, and ESS alignment done. c. then BSS feedback told us readings from P1E and P2E were out of tolerance. During ESS alignment, it used 15 nA, which was considered too big. Tried to change this parameter in config to 5 nA, but don’t have authorization. 2. MID 3242. a. Firstly we followed @Zheng Yang’s suggestion to check if data was the same as yesterday after we re-setup the water phantom. There’s a error about 0.7 mm. then we setup and save a new isocenter again, the error decreased to 0.3mm. after discussion, the solution is, save the photo of the coordinates of isocenter and trust this value, then use a gauge block to check if it was good enough, considering the mechanical error of the motor. b. Then, we scanned 4 curves to see if different direction of the detector will give variable values of R80 (as I mentioned yesterday with @Jake Hindenburg). The conclusion was, there’re differences, but not the same between two curves. But we’ll always use the same scanning direction (distal to proximal). c. Finally everything was good to go. During measurement of B6, we lost LLRF. IMPA and FPA was shut down because of interlocks. Reset SRCU and clear alarms again and again, then we had RF back. d. When trying to find the corresponding range of the lower edge of B5, we were blocked by the validation of B1234E current. Tried several method: i. Manually find the best B1234E current (whick can let P3E has highest reading) according to the degrader motor step, and retrodicted the correct set range, then iterate. ii. Changed degrader motor step (original: 11950. We tried 11900 and 11980), then iterate. These two method will lead us back to the failed point after iteration. iii. Then we found the former iteration process went from lower range to higher ones. So we tried to give it a much higher setpoint of range to iterate. Then, we do get a new correct set range, but the new B1234E current still cannot pass the validation. e. So we had to go back and check the of MID3252. Then we found there’s one point, which as a error of 10 gauss (tolerance is 10 gauss) of the NMR reading between the old value and the value we acquired just now. This error caused the polynomial and fitted curve changed a lot. So we need to update ConvAlgo with new B1234E parameter and do MID3242 again. | Xinyao Li | 06/28/20 8:35 PM
Row 115 | By Mike: 18 June 2020 4. Before dinner, the MID 3242 started with film validation between block 3 and 4 to check the degrader location in the beam position. The EBT film image is in public drive. 5. Moving on the degrader calibration measuring, we started from the B9 (range 33.43~31.6). The process is quite smooth, we can get the acceptable result in two runs. Later, Passthrough position, B8 (range 31.73~29.82) and the largest in B7 (range 30) were completed. | Xinyao Li | 06/28/20 8:35 PM
Row 115 | redo all blocks, test and pass | Chenchen Yuan | 06/28/20 8:12 PM
Row 115 | finish block 9 and block 8 test | Chenchen Yuan | 06/19/20 12:02 PM
Row 115 | motor step point finished | Yeehim Him | 05/28/20 11:00 PM
Row 116 | Reviewed and OK for CCB T30. @janet.huang@iba-group.com , please help scan and upload to the corresponding bug. | Sherman Chang | 07/03/20 3:39 PM
Row 116 | for now, Because of SL1E motor broken, automatic mode can't use, every day shift we need manually alignment beam on ESS axis; | Chenchen Yuan | 07/01/20 3:56 PM
Row 116 | paper work finish, and send to Mikhail for reviewing | Chenchen Yuan | 07/01/20 3:55 PM
Row 116 | finished procedure except paper work and update bugzilla | Yeehim Him | 05/26/20 11:54 PM
Row 117 | Reviewed and OK for CCB T30. @janet.huang@iba-group.com , please help scan and upload to the corresponding bug. | Sherman Chang | 07/03/20 3:39 PM
Row 117 | test result pass; paper work finish, and send to Mikhail for reviewing | Chenchen Yuan | 07/01/20 3:54 PM
Row 117 | With Chenchen, Douglas and Mike Summary: redid MID 3252 and finished. Details: 1. As discussed before, we checked some spots in MID3252 and decided to redo it. 2. Stopped by unstable RF several times. 3. Also sometimes we cannot load current to B1234E (there’s reading from P3E, but the current was lower than normal). Tried to turn off the power to B1234E then active the magnet again, but to no avail. Then tried to reload set BLAK files, everything was fine (need to do it several times, load different set BLAK files to “trigger” the system….). We guess it’s caused by the communication between BLAK and BCP/achestra. 4. Step of B1234E was changed from 0.75 A to 0.4 A after set range lower than 22 g/cm2. We found the fitting curve was better to find a more precise value of max current for B1234E. Below is a capture comparing step of 0.75A and 0.4A for lower range. Finally we finished MID3252 and the fitting curve looks better this time. | Xinyao Li | 06/22/20 11:46 PM
Row 117 | finish all test and passed | Chenchen Yuan | 06/19/20 12:01 PM
Row 117 | finished all range test according to MID 3252, but some ranges need validation again; detail report as below: 1. Started work and found beam size was too big on P1S1 and cannot be focused adjusting Q1S1 and Q2S1. Then checked beam line and realized the vacuum was not good because of the power outage. We turned on valves and wait for it became good enough to continue work. 2. Then we continued work and found current was too low on P3E. we checked the devices in beam line from the exit of cyclo to WP and found the setpoints of slits were incorrect. Changed the setpoints of all 3 slits and beam profile was good to go. 3. During measurements, RF was still unstable and stopped us several times. Need to clear alarms ,reset SRCU, and wait until it recovered. Painful. 4. As @Mikhail Badylevich required, we redo the crossline and depth scanning before dinner, and the capture of curves are in shared drive. 5. After dinner, we cannot increase the current of Q1S1. We power cycle the controller of MPSS, then restarted BLCU through PLC computer. Then quads were good. 6. Because BLCU was restarted, we turned on source and redo mail coil. RF was still painful. 7. After system recovered, SL3E went wrong during we tried to load setpoints to the devices. We reset MSEU, SL3E was good from Achestra, but still cannot receive correct position info. Restarted MSEU again, and we can home SL3E successfully. 8. It’s a little bit tricky for the lowest range. a. For the last point (range = 4 g/cm2), we got R80 = 4.177 (larger than 4) (motor step 21398) from scanning of WP. b. So we decided to use a lower value of range (3.8), then it’s possible to calculate a average value closer to 4. The correct R80 for 3.8 g/cm2 is 4.015 g/cm2 (motor step 21473). c. After that, we rescale motor step to 21480 ( = (21473-21398)/(4.177-4.015)*0.015+21473), and it passed the test. 9. The errors of gauss of three points were out of tolerance, and the field NMR of one points seems incorrect. But system was really unstable when we finished the last measurement. So we decided to validate them on next shift. 10. The highest and lowest range are 33.5 and 4 respectively. | Chenchen Yuan | 06/18/20 1:05 PM
Row 117 | With Chishuo, Chenchen and Douglas. Summary: almost finished MID3252. Data and captures of scanning curves saved in shared drive. Details: 1. Started work and found beam size was too big on P1S1 and cannot be focused adjusting Q1S1 and Q2S1. Then checked beam line and realized the vacuum was not good because of the power outage. We turned on valves and wait for it became good enough to continue work. 2. Then we continued work and found current was too low on P3E. we checked the devices in beam line from the exit of cyclo to WP and found the setpoints of slits were incorrect. Changed the setpoints of all 3 slits and beam profile was good to go. 3. During measurements, RF was still unstable and stopped us several times. Need to clear alarms ,reset SRCU, and wait until it recovered. Painful. 4. As @Mikhail Badylevich required, we redo the crossline and depth scanning before dinner, and the capture of curves are in shared drive. 5. After dinner, we cannot increase the current of Q1S1. We power cycle the controller of MPSS, then restarted BLCU through PLC computer. Then quads were good. 6. Because BLCU was restarted, we turned on source and redo mail coil. RF was still painful. 7. After system recovered, SL3E went wrong during we tried to load setpoints to the devices. We reset MSEU, SL3E was good from Achestra, but still cannot receive correct position info. Restarted MSEU again, and we can home SL3E successfully. 8. It’s a little bit tricky for the lowest range. a. For the last point (range = 4 g/cm2), we got R80 = 4.177 (larger than 4) (motor step 21398) from scanning of WP. b. So we decided to use a lower value of range (3.8), then it’s possible to calculate a average value closer to 4. The correct R80 for 3.8 g/cm2 is 4.015 g/cm2 (motor step 21473). c. After that, we rescale motor step to 21480 ( = (21473-21398)/(4.177-4.015)*0.015+21473), and it passed the test. 9. The errors of gauss of three points were out of tolerance, and the field NMR of one points seems incorrect. But system was really unstable when we finished the last measurement. So we decided to validate them on next shift. 10. The highest and lowest range are 33.5 and 4 respectively. | Xinyao Li | 06/18/20 1:04 PM
Row 117 | With Chenchen, Douglas and Yeehim Summary: Finished MID3252 8/15, and the first 3 need to be validated for not focused well before WP. Details: 1. Continued work left from yesterday, cyclo was good. 2. 15 ranges in MID3252 need to be tuned, and we finished 8 of them. It should be mentioned that, after we finished tuning the first 3 ranges, we found the beam was not focused on PIS1 (the last BPM before WP) on Y axis, the width was over 25mm. 3. Then we thought Transport need to be used. But after discussing with Hassan, we decided to tune Q1S1 and Q2S1 manually to get better focused beam on P1S1. 4. Since beam size is small on high ranges, we moved on to the 4th range (the first 3 ranges will be validated at last). 5. Things went smoothly until 20:40, and RF became really unstable. Since we need to prepare work for the power cycle tomorrow, we decided to save data and that’s the end of the shift. 6. Data saved in shared drive (set BLAK files, WP curves and MID3252 spreadsheet). Tips: During using WP, we found the detector should move from distal to proximal to get better data. We tested 2 routes, distal to proximal, and proximal to distal, and there’s a difference of 0.02 cm, and this tiny difference caused failure during one test. | Xinyao Li | 06/14/20 9:03 PM
Row 118 | Reviewed and OK for CCB T30. @janet.huang@iba-group.com , please help scan and upload to the corresponding bug. | Sherman Chang | 07/03/20 3:39 PM
Row 118 | test result pass; paper work finish, and send to Mikhail for reviewing; | Chenchen Yuan | 07/01/20 3:53 PM
Row 118 | validate BLUE phantom systematic error; recheck 4 range: pass through, high, middle, low: error lower than 0.1mm | Chenchen Yuan | 07/01/20 3:52 PM
Row 118 | use new Field to current fitting parameters (B1234E) to update ESBTS convalgo; redo range 33.1, 32.5~4 cm; BLUE phantom systematic error; ask help from expert, | Chenchen Yuan | 07/01/20 3:51 PM
Row 118 | first round finished, need expert review the result, to confirm if we need recheck or not | Chenchen Yuan | 06/28/20 8:16 PM
Row 121 | need create relate bug | Chenchen Yuan | 07/21/20 3:37 PM
Row 121 | Add MID 65063 in ML5a. From Nicolas Gerard : R2VCU system ( redundant range verifier control unit ) has been added on Proteus Plus to be compliant to 60601-2-64. | David Ringenbach | 12/21/18 8:17 PM
Row 135 | create related bug in bugzilla | Chenchen Yuan | 07/21/20 3:39 PM
Row 135 | Add MID 61281 in ML10a. From Nicolas Gerard : R2VCU system ( redundant range verifier control unit ) has been added on Proteus Plus to be compliant to 60601-2-64. | David Ringenbach | 12/21/18 8:19 PM
Row 138 | @Yangzhao, please create related bug in bugzilla | Chenchen Yuan | 07/21/20 3:40 PM
Row 138 | Add MID 73288 in ML10c. "From Ingo Seidel : The PRIDEx cloud connector has to be installed on every site. MID73288 is describing the full installation and set up of the required virtual machines and required software components. MID73288 can be executed as soon the PTS software is running. As it is not room specific it can be done quite early in the process. MID47147 (already part of the GSTM) is required to validated the cloud connector. " | David Ringenbach | 12/13/18 8:29 PM
Row 156 | @Jianlong, please create related bug in bugzilla | Chenchen Yuan | 07/21/20 3:40 PM
Row 156 | Add MID 68811 in ML6a. From Paul-Francois Doubliez : Correction of LEONI A6 offset for visual alignment of the couch relatively to the lasers. BUG 88027 | David Ringenbach | 12/21/18 7:48 PM
Row 157 | The test is failed in step 6.1.2, inside the limit switch some wiring connect wrongly, there is also a same issues in room3, so will update the limit box wiring later and copy it from room3. | Jun Zhao | 07/01/20 7:40 PM
Row 160 | Fnish the TPO ,everything is ok. | Jun Zhao | 07/01/20 7:29 PM
Row 161 | 1.Finish the test except the 7.4.2 and 7.4.4, and as the pcu is not available on site, these two steps need apply 24V, so will test later. 2 .Also there is some issues with the step 7.4.5 in room3, so we plan check it again also discuss with Jianlong. | Jun Zhao | 07/01/20 7:34 PM
Row 167 | add one cable from Bopla TSS box in DNSHCU to SRCU. | Jianlong Wen | 06/19/20 2:57 PM
Row 232 | Remove MID 6712 from ML9b. Florence Kremer : BUG 100826 | David Ringenbach | 12/12/18 10:45 PM
Row 288 | @Jianlong, please create related bug in bugzilla | Chenchen Yuan | 07/21/20 3:41 PM
Row 288 | Add MID 68811 in ML6a. From Paul-Francois Doubliez : Correction of LEONI A6 offset for visual alignment of the couch relatively to the lasers. BUG 88027 | David Ringenbach | 12/21/18 7:49 PM
Row 363 | Remove MID 6712 from ML9b. From Florence Kremer : Bug 100826 | David Ringenbach | 12/12/18 10:46 PM
Row 407 | @shuling.peng@iba-group.com Please update status | Sherman Chang | 07/20/20 9:52 AM
| Since the SRCU3 is not connected to the SRCU0, the interconnection test can not be carried at the moment. | Shuling Peng | 07/22/20 10:03 AM
Row 407 | April 14th, 2020 Done with TCR test | Shuling Peng | 04/14/20 4:15 PM
Row 407 | April 13th,2020 cable installation for relay box was done, and signal test for relay box has passed | Shuling Peng | 04/13/20 4:06 PM
Row 407 | Jun 13th, 2020. With JEFF. We found a discrepancy fault caused by missing HPV3 wmebcu. The problem cannot be solved until the HPV3 is installed. | Shuling Peng | 01/14/20 8:59 AM
Row 407 | Jun 10th, 2020. Done with emergency bottom debug | Shuling Peng | 01/14/20 8:46 AM
Row 407 | Jun 7th, 2020. All emergency bottoms except for two are now OK. The entrance door limit switch is also OK now. The two abnormal ones need to contact with control team to check the code. | Shuling Peng | 01/07/20 9:32 AM
Row 407 | DEC 19th,2019. Done with the debug for SRCU3 door limit switches. | Shuling Peng | 12/19/19 5:07 PM
Row 407 | DEC 17th,2019. Done with bypass for signal signs and door limit switch which were not installed yet in the gantry pit. But some debug work still needs to be done. | Shuling Peng | 12/18/19 8:20 AM
Row 407 | DEC 16th,2019. The Phase One for the signal test has been done. Since only 6 out of 10 ES and 2 out of 3 SB were installed, I have to bypass other signal signs first to make the Gantry pit search procedure work so that other other procedures can be carried on. | Shuling Peng | 12/16/19 5:11 PM
Row 408 | Reviewed by CCB T30: xwchang, ccyuan. OK to close and set to QA for validation. | Sherman Chang | 07/22/20 10:18 AM
Row 408 | 2020-7-7 Done with MID 61545 for SRCU0 & SRCU3, and handed in the paperwork | Shuling Peng | 07/07/20 4:42 PM
Row 411 | Reviewed by CCB T30: xwchang, ccyuan. OK to close and set to QA for validation. | Sherman Chang | 07/22/20 10:18 AM
Row 411 | Telescope is separated in 5 part, and each part need test 5 different range (32.5, 24.5, 15.5, 10.5, 4.0). we used Blak, Convalgo and transport. We finished 2 part (straight line 1 & 2), result was reviewed by expert, and passed, need paper work! | Chenchen Yuan | 07/08/20 1:40 PM
Row 412 | update new Table generator | Chenchen Yuan | 07/08/20 1:44 PM
Row 412 | ESBTS convalgo (all ess calibration parts )had updated, and add telescope parts, upload in public drive | Chenchen Yuan | 07/08/20 1:43 PM
Row 419 | @Jianlong, please create related bug in bugzilla | Chenchen Yuan | 07/21/20 3:41 PM
Row 419 | Add MID 68811 in ML6a. From Paul-Francois Doubliez : Correction of LEONI A6 offset for visual alignment of the couch relatively to the lasers. BUG 88027 | David Ringenbach | 12/21/18 7:50 PM
Row 420 | we have finish the procedure, only need to do the paper work. | Jianlong Wen | 07/01/20 5:52 PM
Row 421 | work with Xiaogang, finish 35%. | Jianlong Wen | 06/11/20 3:55 PM
Row 422 | have done the procedure, only need to do the paper work. | Jianlong Wen | 07/01/20 5:48 PM
Row 423 | have done the procedure, need to do the paper work. | Jianlong Wen | 07/01/20 5:41 PM
Row 425 | @deyou.wang@iba-group.com Please add comments also while you make any progress. Thank you! | Sherman Chang | 07/22/20 10:21 AM
Row 426 | @deyou.wang@iba-group.com Please add comments also while you make any progress. Thank you! | Sherman Chang | 07/22/20 10:21 AM
Row 427 | @deyou.wang@iba-group.com Please add comments also while you make any progress. Thank you! | Sherman Chang | 07/22/20 10:21 AM
Row 428 | @deyou.wang@iba-group.com Please add comments also while you make any progress. Thank you! | Sherman Chang | 07/22/20 10:22 AM
Row 429 | finish all the procedure except burndy1-5, in step 009 and burndy6-1, in step057. for the first one, becuse customer didn't finish the installation in GTR3, for the second one, need clarify with experts. | Jianlong Wen | 07/17/20 2:52 PM
Row 429 | Some step is blocked due to the hpv3 not finish the installation , now it is working normally ,will do the test after gefos mission. | Jun Zhao | 07/01/20 7:58 PM
Row 430 | finish the step [DNSH-ITC-015], all the data is meet the required in TPO. | Jun Zhao | 07/01/20 7:44 PM
Row 430 | During doing the procedure, meet the problem, the DNSHCU can not know the presence of the tray, and when insert the tray, the error occur: ACCESSORIES TRAY:Discrepancy on locking. | Jianlong Wen | 06/10/20 4:37 PM
Row 430 | clean the TCR3 and help to resume the system. after that we do the DNSH IT. | Jianlong Wen | 06/08/20 3:58 PM
Row 430 | since last thurday, together with jun,we are playing MID 45578 SRCU DNSH IT,about 70% finshed, MID37826 just started, MID 45778 SRCU GTR IT ,about 65%finished, MID 43441 SRCU PPS IT, about 70%finshed; left items are related to FOV. Today we lost almost the whole shift,because when we start the morning shift,the process seems back to two or three weeks ago, DNSH sew down,modbus is down,HPV3 locked,when we retsart DNSH to reset SEW,but modbus is still down,WMEBCU down,we ask jianlong and Yangzhao for help,when Yangzhao try to unlock HPV3,BPMEU and DNSH tripped from PSR.After power on,HPV3 cannot be locked,not sure due to register or other problem.Fortunately at last after jianlong register muanally,HPV3 unlocked.For DNSH,we know there is a bug there,but other problems,we not sure why and we see ' ok' from last shift report. | Deyou Wang | 05/14/20 4:34 PM
Row 432 | it has been done by Leoni, need to upload the report. | Jianlong Wen | 04/15/20 1:53 PM
Row 433 | it has been done by Leoni, need to upload the report. | Jianlong Wen | 04/15/20 1:54 PM
Row 434 | Work with Jianlong 1. Do the preparation for the coupler installation, then run the TPO. | Jun Zhao | 05/25/20 7:01 PM
Row 435 | we have finish the procedure, only need to do the paper work. | Jianlong Wen | 07/01/20 5:55 PM
Row 435 | help BETTER to install the counterweight on the GTR3. | Jianlong Wen | 06/08/20 3:59 PM
Row 435 | Work with Xiaogang 1. Do the gantry rough balancing, there is stall happened when CW the gantry at 310 degree no moving, then add 300kg on the counter weight 1.(facing the nozzle), test again, no stall. | Jun Zhao | 05/26/20 5:23 PM
Row 442 | 2020/3/09~2020/3/11 With Jianlong, done with MID-2017 except for belt tension test due to the lack of equipment. The belt tension test is to be done after the equipment is purchased | Shuling Peng | 03/11/20 5:11 PM
| belt tension work is done, all the TPO is finished except the paper work | Jun Zhao | 07/01/20 7:50 PM
Row 443 | Finish all the test except the paper work. | Jun Zhao | 07/01/20 7:50 PM
Row 445 | finish the rest of the procedure, need to finish the paper work. | Jianlong Wen | 07/15/20 4:31 PM
Row 445 | The function is ok, as the gantry is moving now, we will do the rest after gefos mission | Jun Zhao | 07/01/20 7:55 PM
Row 446 | @deyou.wang@iba-group.com Please add comments also while you make any progress. Thank you! | Sherman Chang | 07/22/20 10:22 AM
Row 448 | The function of LTEU is ok, anyway as the CBCT workflow is not start normally, we define the protective field much small than the PTO asked. as we are in the installation phase, will resume the normal state when the CBCT is totally ready. | Jun Zhao | 07/01/20 7:48 PM
Row 448 | work with Jun: 1.Trace the root cause with of the LTEU can not power on, always shorted even disconnect the plug on the Rotoscan, we check the cable find the cable is not weld well after reweld it works 2. configure the two Rotoscan, copy the config data to Guangzhou PMS folder. 3. Uninstall the laser rack from the PCU cabinet, print the drawing to prepare the voltage converting. | Deyou Wang | 03/30/20 10:55 PM
Row 449 | @deyou.wang@iba-group.com Please add comments also while you make any progress. Thank you! | Sherman Chang | 07/22/20 10:22 AM
Row 451 | @deyou.wang@iba-group.com Please add comments also while you make any progress. Thank you! | Sherman Chang | 07/22/20 10:22 AM
Row 453 | @deyou.wang@iba-group.com Please add comments also while you make any progress. Thank you! | Sherman Chang | 07/22/20 10:22 AM
Row 454 | Worked with Jianlong | Xiaogang Ji | 06/04/20 10:39 PM
Row 454 | 1. Finally, we can acquire image through AI for radA and radB. Thanks again for Loic’s support. For radB exposure issue, it relate to the bug83737, solution was provided by Loic: a. Connect the cable 90W3 under the cable 90W4 in the back of the PPVS computer; b. Swap the cable 90W3 and 90W4 in the PLC; After this, power on the PPVS cabinet and PPVS computer again. Reset the error, radB can acquire image through AI. This modification has been apply to GTR1 and GTR2 as well. For radA exposure issue, my colleagues power cycle the PPVS cabinet and EMD, and we try again to shoot x-ray in AI, and it success. For the ‘’flat panel interlock’’ in the Generator1, it’s still flashing when we didn’t connect to AI, but while the XCU connect to AI and we load the patient for shooting, the ‘’flat panel interlock’’ is stop flashing, and we can shoot the x-ray. | Jianlong Wen | 06/04/20 3:07 PM
Row 454 | work with Jianlong Summary 1. The config file for EMD has been updated, so we need to upload the EMD file again, and redo the tube calibration. (MID54462) 2. Acquisition of X-ray images with service software（MID40485）for Calibrating the Varian Paxscan 4030 flat panel, but x-ray(tube A) exposures failed (The “Host:port” = 10.2.53.73:502 ) 3. Troubleshooting for the X-ray exposure. Issues： 1. Loading the CBCT workflow in the PPVS PC, the AdaPT insight crash. 2. Tube A can’t expose. | Xiaogang Ji | 06/03/20 9:55 PM
Row 454 | the config file for EMD has been updated, so we need to upload the EMD file again, and redo the tube calibration. | Jianlong Wen | 06/03/20 3:20 PM
Row 456 | fill in the data in this procedure. | Jianlong Wen | 07/16/20 4:32 PM
Row 456 | fill in the data in test report. finish flat panel A calibration, but the image in AI is not good, this problem need help from experts, will ask for help from the experts available tomorrow. | Jianlong Wen | 07/15/20 4:31 PM
Row 456 | the installation plate for laser3 is upside down, it takes us about 6 hrs to fix this problem. no matter what, laser1,2,3,4 are aligned roughly base on the cross wire on pps test couch. | Jianlong Wen | 07/14/20 4:20 PM
Row 456 | because we didn't use the latest ai-XCU client, we can not set run length, it's hard to catch image using TTT. it takes us about 5.5hrs to find out why we can not shoot continuous X-ray in CBCT mode. this is very precious experience for everyone. for the rest of the shift, we try to align the PPS test couch. | Jianlong Wen | 07/06/20 10:47 PM
Row 456 | finish the flat panel alignment for radA and radB. | Jianlong Wen | 07/03/20 11:03 PM
Row 456 | Finally, we have the tool 577120, and can continue the Gefos mission. we spend about 4hrs to do the preparation, and spend 4hrs to proceed the procedure that Gefos send to us. So far so good. | Jianlong Wen | 06/30/20 10:06 PM
Row 456 | Leica need the tool 577120 to do the mission, can not start the Gefos mission today, try to do the FPB calibration, but the PU is down, spend about 3 hrs to recover the system. | Jianlong Wen | 06/29/20 11:10 PM
Row 456 | preparation for the PPS universal tool for the gefos mission. | Jianlong Wen | 06/24/20 6:05 PM
Row 456 | 1. Go through the MID40485 for radA FP calibration; 2. Try to run MID40496 for radB FP calibration, something is not clear now, afternoon shift can continue to do it; | Jianlong Wen | 06/18/20 3:24 PM
Row 457 | @deyou.wang@iba-group.com Please add comments also while you make any progress. Thank you! | Sherman Chang | 07/22/20 10:22 AM
Row 458 | finally, we found out the root cause why radA is fail to exposure by AI. there are wrong connection about the J14 in the EMD site. | Jianlong Wen | 06/12/20 4:00 PM
Row 467 | After load the config file from Loic, we can run CBCT now, Yang will give the details info about this. | Jianlong Wen | 06/18/20 3:24 PM
Row 467 | we want to do the CBCT, but A.I. crash again. still need to find out why. | Jianlong Wen | 06/16/20 5:50 PM
Row 469 | @deyou.wang@iba-group.com Please add comments also while you make any progress. Thank you! | Sherman Chang | 07/22/20 10:22 AM
Row 471 | And what's the meaning of 1% ? | Sherman Chang | 07/22/20 10:22 AM
Row 471 | @deyou.wang@iba-group.com Please add comments also while you make any progress. Thank you! | Sherman Chang | 07/22/20 10:22 AM
Row 475 | Reviewed and OK for CCB T30. @janet.huang@iba-group.com , please help scan and upload to the corresponding bug. | Sherman Chang | 07/03/20 3:38 PM
Row 475 | LPIC is ok now | Chishuo Chuang | 06/09/20 8:42 PM
Row 475 | start vacuum of GTR3, according to ZZ site suggestion, we need use AS4.2 to force active the flush control, this step can trig LPIC to "safe mode", then wait vacuum drop down to 18 mbar, click "reset" button on VNC, "fill" valve will open, rise the IC1 inside vacuum to 200 mbar, status switch to "ready to beam" | Chenchen Yuan | 06/04/20 11:19 PM
Row 475 | we installed the LPIC1 in GTR3. Except the final step of Vacuum test, other tests from the MID is passed. | Yeehim Him | 06/04/20 12:07 AM
Row 480 | SMPS cabinet need replace. waiting shipment from LLN | Yeehim Him | 07/21/20 11:58 PM
Row 480 | start SMPS IT, SMPSEU configuration is finished; | Chenchen Yuan | 06/04/20 11:26 PM
Row 495 | Remove MID 6712 from ML9b. From Florence Kremer : Bug 100826 | David Ringenbach | 12/12/18 10:47 PM
Row 511 | Guys. Please add comments also while you make any progress. And what's the meaning of 1%Thank you! | Sherman Chang | 07/22/20 10:23 AM
Row 514 | after power on the FOVCU and home the blades, the left and right blade can not move, check the cable, no problem, restart the cabinet and home the blades again, the left and right blade can be moved now. The root cause is not clear. | Jianlong Wen | 06/16/20 5:47 PM
Row 514 | with jun,check the cable WB3421 and reconnect the burndy, then it works.HPv3 locked,then register manually and turn the FOVcu and DNSHcu to managed mode,then unlocked. Try to test rpd and virtual HP, but it needs to home imager firstly,for now homing always comes an error from drive FOV.still need to check. | Deyou Wang | 05/15/20 4:31 PM
Row 537 | This is the last line of the Site Test Matrix. Do ALL work above this line. | BJones 1-USA | 01/26/17 5:11 AM
Sheet | work with Jun, 1.Make the Power Supply cable and plug for WMEBCU and the UPS on the wall in pulling room. ,2.Make the Burdy of the two big cable WB3273 and WB3276 both to WMEBCU and test the continuity of each pin. 3,replace two existing old network cable from the PCU to network switch by CAT 6, test the new cable is good. 4.replace the old network cable from the WMEBCU to Network switch by CAT 6, cable tested is good. but not connected. | Deyou Wang | 04/01/20 10:52 PM
Sheet | 2020/3/19-2020/3/20 (1) GTR3: with Jianlong, working on MID-2521 test procedure (2)SRCU: with Zhao Yang, debug PLC laptop connectivity issue | Shuling Peng | 03/20/20 5:01 PM
Sheet | 2020/3/18 (1) GTR3: with Jianlong, working on MID-2515 test procedure in the morining (2)SRCU: with Zhao Yang, Version Dog upgrade on server in the afternoon Ps: since the customer contractor still not onsite, the signal signs installation for SRCUi cannot be proceed now | Shuling Peng | 03/18/20 4:54 PM
Sheet | 2020/3/17 (1) GTR3: with Jianlong, working on MID-2019 test procedure in the morining (2)SRCU: Version Dog upgrade & SRCU3 rollback in the afternoon Ps: since the customer contractor still not onsite, the signal signs installation for SRCUi cannot be proceed now | Shuling Peng | 03/17/20 3:06 PM
Sheet | 2020/3/16 (1)SRCU3: a new code SRCU-C3.0a needs to be deplyed for SRCU3, but the vDOG needs to be upgraded before the new code deployment. Yang is working on the upgrade test, waiting for feedback. (2) GTR3: with Jianlong, working on MID-2019 test procedure. | Shuling Peng | 03/16/20 4:06 PM
Sheet | 2020/3/13 Deploy upgrade for SRCU0 & SRCU3 | Shuling Peng | 03/16/20 11:28 AM
Sheet | Nov 2019, Tus, on shift with Jeff. We tested the SCRU0 again. But the system seemed not stable. Sometimes it worked all right, sometimes it occoured an short-circuit alarm. We assumed it's the problem with the customer relays. We have requested the customer to change the relays. | Shuling Peng | 11/26/19 4:01 PM
Sheet | put back the Dee, start vacuum. Test the H3 with compress air, vacuum still rise from 4.6E-6 to 6.7 E-6. So we confim the leak is from the Dee | Lipeng Han | 10/28/19 10:00 AM